## Ext LDT
Name of Test (linked to EUA summary) | Laboratory (linked to EUA letter) | Type of test | Sensitivity | Specificity | Specimens needed for the test | Target genes | Limit of detection (LoD) | Test details | Where/how is it being used now?
Quest PF COVID-19 PCR DTC | Quest Diagnostics Nichols Institute | rRT-PCR | 100% | 100% | anterior nasal swabs | ORF1ab gene | 1 x 10^-2 TCID50/mL | The Quest PF COVID-19 PCR DTC is a direct to consumer product intended for the qualitative detection of nucleic acids from SARS-CoV-2 in anterior nasal swab specimens collected using the Quest COVID-19 PCR Test Home Collection Kit. The test is also intended for the qualitative detection of nucleic acids from SARS-CoV-2 in pooled samples containing up to five individual anterior nasal swab specimens collected using the Quest COVID-19 PCR Test Home Collection Kit. The Hologic Panther Fusion SARS-CoV-2 Assay provides the positive and negative controls with the kit, and the controls are ready-to-use. The specimens collected with the Quest COVID-19 PCR Test Home Collection Kit are tested with the Hologic PantherFusion SARS-CoV-2 Assay on the Panther Fusion System. | Testing is limited to laboratories designated by Quest Diagnostics, which are also certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meet the requirements to perform high-complexity tests.
Quest RC COVID-19 PCR DTC | Quest Diagnostics Nichols Institute | rRT-PCR | 100% | 100% | anterior nasal swabs | ORF1ab, E genes | 0.009 TCID50/mL | The Quest RC COVID-19 PCR DTC is a direct to consumer product intended for the qualitative detection of nucleic acids from SARS-CoV-2 in anterior nasal swab specimens collected using the Quest COVID-19 PCR Test Home Collection Kit. The test is also intended for the qualitative detection of nucleic acids from SARS-CoV-2 in pooled samples containing up to six individual anterior nasal swab specimens collected using the Quest COVID-19 PCR Test Home Collection Kit. The test is performed by testing anterior nasal swab specimens collected with the Quest COVID-19 PCR Test Home Collection Kit with the Roche cobas SARSCoV-2 assay on the Roche cobas 6800/800 Systems. | Testing is limited to laboratories designated by Quest Diagnostics which are also certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meet the requirements to perform high-complexity tests.
Quest COVID-19 PCR Test Home Collection Kit | Quest Diagnostics Nichols Institute | Home collection | N/A | N/A | anterior nasal swabs | N/A | N/A | The Quest COVID-19 PCR Test Home Collection Kit is a direct to consumer (DTC) product for collecting an anterior nasal swab (nasal) specimen at home (which includes in a community-based setting), from individuals age 18 years and older (self-collected), 16 years and older (self-collected under adult supervision), or 2 years and older (collected with adult assistance) that is sent for testing with an in vitro diagnostic (IVD) molecular test that is indicated for use with the Quest COVID-19 PCR Test Home Collection Kit. All test results are delivered to the user via an online portal. Individuals with positive or inconclusive/invalid results will be contacted by a healthcare provider. | Testing is limited to laboratories designated by Quest Diagnostics that are certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meet the requirements to perform high complexity tests.
Amazon Real-Time RT-PCR Test for Detecting SARS-CoV-2 | STS Lab Holdco (a subsidiary of Amazon.com Services LLC) | Home collection, rRT-PCR, pooling | 95.7% | 100% | Anterior nasal swabs | ORF1ab | 500 copies/mL | The January 2022 revision updated the test's standard operating procedure. The December 2021 revision added alternative packing companies for collection kit assembly and alternative suppliers for collection kit components. The Amazon Real-Time RT-PCR Test for Detecting SARS-CoV-2 should be used with the Amazon COVID-19 Collection Kit. The test can be used in individuals who are asymptomatic, and who are not suspected of COVID-19 by their healthcare provider. The test is by prescription use only. Sample collection can be under the supervision of a healthcare provider, or unsupervised at home (in individuals 18 and older). Samples may also be pooled with up to 5 anterior nasal swab samples per pool. Amazon is planning to use this in their employee testing plan. The rRT-PCR assay is a modification of the BGI Genomics Real-Time RT-PCR Test. Usability studies were performed in individuals 18 and older. Sensitivity was determined using 69 positive clinical samples from asymptomatic patients, and specificity was determined from 270 negative clinical samples from asymptomatic patients. | The test can be run by CLIA laboratories designated by STS Lab Holdco (a subsidiary of Amazon.com Services LLC). Amazon is currently using the test for employee screening.
Color SARS-CoV-2 RT-LAMP Diagnostic Assay | Color Health, Inc. | RT-LAMP | 100% | 100% | nasopharyngeal swabs, oropharyngeal swabs, anterior nares swabs (also via self-collection), mid-turbinate nasal swabs, nasopharyngeal wash/aspirate or nasal aspirates, as well as bronchoalveolar lavage specimen | N, E genes | 750 copies/mL | The January 2022 revision updated Color COVID-19 Self-Swab Collection Kit DTC to include additional usability study data on 2-4 year-olds using modified collection instructions and provide minor edits. The July 2021 update includes an update to the authorized labeling to change the assay format to include an initial screening assay followed by reflex testing for invalid results and candidate positive specimens. The April 2021 update includes an update to sample collection: the test may now be used for asymptomatic individuals, and self-collection using the Color COVID-19 Self-Swab Collection Kit or the Color COVID-19 Self-Swab Collection Kit with Saline. This test also has a version that can be marketed as a direct-to-consumer product as of 3/19/21. The August and November updates include the removal of the ORF1ab primer sets from the protocol, updated assay and control procedures, a name change to the "Color Genomics SARS-CoV-2 RT-LAMP Diagnostic Assay," and clarifying language for the protocol. The IFU was also updated with the NIH reference panel performance data. This assay uses reverse transcription and loop-mediated isothermal amplification (LAMP) to provide a rapid, qualitative colorimetric readout. RNA extraction is to be performed using the Chemagic 360 platform. RNA samples can then be run on a 384 well plate, using the Hamilton STARlet system 4 and the Biotek NEO2 microplate reader. The test takes 70 minutes. There is also an internal control, human target gene (RP). Sensitivity and specificity were determined using contrived and clinical samples, though only clinical sample data is presented here (41 positive, 502 negative). The LoD according to the NIH reference panel is 18,000 NDU/mL. | Healthcare providers can send clinical specimens for testing to Color Genomics, Inc. (Burlingame, CA)
The Color COVID-19 Self-Swab Collection Kit | Color Genomics, Inc. | Home collection | N/A | N/A | Nasal swabs | N/A | N/A | The January 2022 revision updated the EUA Summary to include additional usability study data on 2-4 year-olds using modified collection instructions and provide minor edits. The EUA was modified from the original issuance on 8/31/2020. The changes include change of the device name from “Color COVID-19 Test Unmonitored Collection Kit” to “Color COVID-19 Self-Swab Collection Kit." It also includes a change in the instructions for use and indication from “unmonitored self-collection” to “unsupervised self-collection." There was also a change in the components of the kit due to supply chain constraints, so that the outer cardboard shipping box was removed, and now it includes a push top closure tube and rigid biohazard bag. The Color COVID-19 Test Self-Swab Collection Kit is coupled with the EUA molecular test by Color Genomics. Swabs from both nostrils are taken, and then placed in prepared tubes and shipped to the Color Genomics lab. This kit does not require doctor supervision via telemedicine. The patient must activate the barcoded kit online before shipping the sample back. The LoD according to the NIH reference panel was 18,000 NDU/mL. | Testing is limited to laboratories designated by Color Genomics Inc. that are certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meet requirements to perform high complexity tests and that run the specimens collected from the Color COVID-19 Test Unmonitored Collection Kit on an IVD molecular test that is indicated for use with the Color COVID-19 Test Unmonitored Collection Kit.
Psoma COVID-19 RT Test | Psomagen, Inc. | rRT-PCR | 100% | 100% | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swab specimens, and bronchoalveolar lavage specimens | N gene (N1 and N2 regions) | 1000 copies/mL | The January 2022 revision updated the authorized labeling of reagent information from "2019-nCoV CDC qPCR Probe assay" to "SARS-CoV-2 Research Use Only qPCR Primer & Probe Kit." The test should be run on the Roche Light cycler 480, with RNA extraction performed separately. Specificity and sensitivity were determined using 55 positive and 55 negative clinical samples. The LoD according to the NIH reference panel is 18,000 NDU/mL. | Testing is limited to Psomagen, Inc
Color COVID-19 Self-Swab Collection Kit with Saline | Color Health, Inc | Home collection, RT-LAMP | 100% | 100% | anterior nasal swab | N, E genes | 750 copies/mL | The January 2022 revision updated the EUA Summary to include additional usability study data on 2-4 year-olds using modified collection instructions and provide minor edits. The Color COVID-19 Self-Swab Collection Kit with Saline should be used in tandem with the Color SARS-CoV-2 RT-LAMP assay. Samples may be self-collected, unsupervised, and in asymptomatic individuals. Currently, this sample collection kit is part of a "community-based distribution framework that is physician ordered. Healthcare providers (HCP) at specific institutions, who are licensed and have prescriptive authority in their respective states, use a COVID-19 eligibility questionnaire that is based on current CDC testing guidelines to evaluate patient eligibility." Color does not accept requests for kits from individuals--this is only for use as part of a community/organization testing plan. The samples are collected and stored in 0.9% saline solution. LOD, Sensitivity and Specificity listed are all reflective of the diagnostic assay from Color Health, Inc. The company also performed multiple stability and usability studies for this sample collection kit. | Testing is limited to CLIA laboratories designated by Color Health Inc.
SARS-CoV-2 MassArray Test | National Jewish Health | rRT-PCR, MALDI-TOF | 100% | 99.5% | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swabs and nasal/nasopharyngeal aspirate and wash specimen, and brochoalveolar lavage | N, ORF1, and ORF1ab | 690 copies/mL | The January 2022 revision updated the EUA summary to include data for the FDA SARS-CoV-2 reference panel testing study, add the Mastercycler Nexus Thermal Cycler realtime PCR instrument for use with the SARS-CoV-2 MassArray Test, update the Limitations section to fulfill Conditions of Authorization related to the September 23, 2021 Viral Mutation Revision Letter, and provide minor edits. The SARS-CoV-2 MassArray Test is a combination of rRT-PCR and mass spectrometry. The rRT-PCR process uses the Agena SARS-CoV-2 Panel Set kit and the Taq-OnePath Master Mix, and should be run on the Thermo Fisher Veriti platform. The amplified sequences are then analyzed using the MassARRAY System for MALDI-TOF. Sensitivity was determined from 39 positive clinical samples, and specificity from 208 negative samples. They also evaluated performance on an alternative platform, Hamilton StarLet liquid handling system, using 32 positive and 74 negative samples. Sensitivity and specificity were both 100%. | Testing is limited to National Jewish Health located at 1400 Jackson St., Denver, CO 80206
PMLS SARS-CoV-2 Assay | Premier Medical Laboratory Services | rRT-PCR | 100% | 100% | anterior nasal swabs | N1, N2 genes | 625 copies/mL | The PMLS SARS-CoV-2 Assay is an in vitro diagnostic real-time reverse transcription polymerase chain reaction (rRT-PCR) Test for the qualitative detection of nucleic acid from SARS-CoV-2 in anterior nasal swab specimens that are collected at home using the INeedaCOVID19Test.com Self Collection Kit. The PMLS SARS-CoV-2 Assay is a modification of the Centers for Disease Control and Prevention (CDC) 2019-novel Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic Panel (EUA200001) and is designed to detect RNA from SARS-CoV-2 in anterior nasal swab specimens. The INeedaCOVID19Test.com Self Collection Kit will be available by prescription to individuals who request Testing for SARS-CoV-2 though eTrueNorth’s on-line platform (ineedacovidTest.com). | Testing is limited to Premier Medical Laboratory Services, 6000A-1 Pelham Rd, Greenville, SC 29615, which is certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263, and meets the requirements to perform high-complexity tests.
Amazon Multi-Target SARS-CoV-2 Real-Time RT-PCR DTC Test | STS Lab Holdco (a subsidiary of Amazon.com Services LLC) | rRT-PCR | 99% | 99.4% | anterior nasal swabs | ORF1ab, N gene | 50 copies/mL | The December 2021 revision added alternative packing companies for collection kit assembly and added alternative suppliers for collection kit components. The Amazon Multi-Target SARS-CoV-2 Real-Time RT-PCR DTC Test (“Amazon Multi-Target DTC Test”) is a direct-to-consumer product for testing individual anterior nasal swab specimens self-collected at home using the Amazon COVID-19 Test Collection Kit DTC by any individual, 18 years or older, including individuals without symptoms or other reasons to suspect COVID-19. All test results are delivered to the user online (via AmazonDx.com), and users receive email and SMS text message notifications that their results are available. The Amazon Multi-Target DTC Test includes primers and probes for the detection of the ORF1ab and N gene regions of the SARS-CoV-2 genome, in addition to human RNase P RNA as an endogenous internal control. | Testing is limited to laboratories designated by STS Lab Holdco (a subsidiary of Amazon.com Services LLC) that are certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meet requirements to perform high complexity tests.
Amazon Real-Time RT-PCR DTC Test for Detecting SARS-CoV-2 | STS Lab Holdco (a subsidiary of Amazon.com Services LLC) | Home collection, rRT-PCR | 95.7% | 100% | anterior nasal swabs | ORF1ab | 500 copies/mL | The December 2021 revision added alternative packing companies for collection kit assembly and alternative suppliers for collection kit components. The Amazon DTC Test is a modification of the BGI Genomics Real-Time Fluorescent RT-PCR Test for Detection SARS-CoV-2 EUA200034). Amazon has obtained a Right of Reference from BGI Genomics Co. Ltd. to information contained in the EUA submission for the BGI Genomics Real-Time Fluorescent RT-PCR Test for Detection SARS-CoV-2 and all current and future amendments. The assay includes primers and probes for the detection of the ORF1ab region of the SARS-CoV-2 genome, in addition to human β-actin RNA as an endogenous internal control. | Testing is limited to laboratories designated by STS Lab Holdco (a subsidiary of Amazon.com Services LLC) that are certified under the Clinical Laboratory Improvement Amendments of 1988 to perform high complexity tests.
Amazon Multi-Target SARS-CoV-2 Real-Time RT-PCR Test | STS Lab Holdco (a subsidiary of Amazon.com Services LLC | rRT-PCR | 99% | 99.4% | anterior nasal swabs | ORF1ab, N genes | 50 copies/mL | The December 2021 update added alternative packing companies for collection kit assembly and added alternative suppliers for collection kit components. The Amazon Multi-Target SARS-CoV-2 Real-Time RT-PCR Test (“Amazon Multi-Target Test”) is an in vitro diagnostic real-time reverse transcription polymerase chain reaction (rRT- PCR) test for the qualitative detection of nucleic acid from SARS-CoV-2 in anterior nasal swab specimens that are self-collected by any individuals (18 years of age or older), including individuals without symptoms or other reasons to suspect COVID-19 (1) onsite at Amazon facilities using either (a) the Amazon COVID-19 Collection Kit under the supervision of a healthcare provider (HCP) or, (b) the Amazon On-Site COVID-19 Test Collection Kit unsupervised or (2) at home using the Amazon COVID-19 Test Collection Kit unsupervised. The Amazon Multi-Target Test is also intended for the qualitative detection of nucleic acid from SARS-CoV-2 in pooled samples containing up to five individual anterior nasal swab specimens per pool that are self-collected in individual vials containing transport medium by any individual. | Testing is limited to laboratories designated by STS Lab Holdco (a subsidiary of Amazon.com Services LLC) that are certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meet requirements to perform high complexity tests.
UCSD EXCITE COVID-19 Test | UCSD BCG EXCITE Lab | RT-PCR | 100% | 100% | anterior nasal swabs | ORF1ab, N, S genes | 1000 copies/mL | The UCSD EXCITE COVID-19 Test is an in vitro diagnostic real-time reverse transcription polymerase chain reaction (RT-PCR) test for the qualitative detection of nucleic acid from SARS-CoV-2 in anterior nasal swab specimens that are either Healthcare Provider collected from individuals of any age, or collected at home. The assay is a real-time reverse transcription polymerase chain reaction (RT-qPCR) test based upon the TaqPath COVID-19 Combo Kit. | Testing is limited to UCSD BCG EXCITE Lab located at 9500 Gilman Drive, San Diego, CA 92161 that is certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meets requirements to perform high complexity tests.
GENETWORx Covid-19 | RCA Laboratory Services LLC dba GENETWORx | Home collection | Not stated | Not stated | nasal swab | Not stated | Not stated | The November 2021 update approved the addition of the Aptima Direct Load Tube Collection Kit as an authorized specimen collection device. This is an at-home collection kit for nasal swab samples for the detection of SARS-CoV-2 when an individual is suspected of COVID-19 by a healthcare provider. The samples are run using the APTIMA SARS-CoV-2 Assay, run in parallel with a GENETWORx RNase P Assay (which serves as a control). The kit may be used unsupervised in individuals 18 and older. Contrived samples were also tested in summer and winter temperatures, with demonstrated stability for up to 36 hours. There is an online ordering system for individuals to purchase kits | Testing is limited to the RCA Laboratory Services LLC located at 4060 Innslake Drive, Glen Allen, VA 23060, and Testing Centers of America LLC located at 411 Swedeland Road, King of Prussia, PA 19406
Labcorp SARS-CoV-2 & Influenza A/B Assay | Laboratory Corporation of America (Labcorp) | rRT-PCR | N/A | N/A | anterior nasal swabs | E gene, ORF1ab (SARS-CoV-2); M1, M2 (influenza A); NEP, NS1 genes (influenza B) | N/A | The Labcorp SARS-CoV-2 & Influenza A/B Assay is a real-time RT-PCR assay intended for the simultaneous qualitative detection and differentiation of SARS-CoV-2, influenza A virus, and/or influenza B virus RNA in anterior nasal swab specimens collected at home by individuals suspected of respiratory viral infection consistent with COVID-19 by a healthcare provider (HCP). The assay is based on the fully automated cobas SARS-CoV-2 & Influenza A/B assay (EUA202635) performed on the cobas 6800/8800 Systems. The anterior nasal swab samples collected with either the Labcorp COVID-19 + Flu Test Home Collection Kit or the Pixel Labcorp COVID-19 + Flu Test Home Collection Kit are stable for up to 56 hours in an ambient temperature shipping environment. | Testing is limited to the Center for Esoteric Testing, Burlington, NC, or other laboratories designated by Labcorp that are also certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meet the requirements to perform high complexity tests.
Life Sciences Testing Center COVID-19 Test | Life Sciences Testing Center | rRT-PCR | N/A | N/A | anterior nasal swabs | N gene, S gene, ORF1ab | 71.4 copies/reaction | The Life Sciences Testing Center COVID-19 Test is a real-time reverse transcription polymerase chain reaction (rRT-PCR) test using the Thermo Fisher TaqPath COVID-19 Combo Kit (EUA200010) with a modified automated extraction method and TaqMan SARS-CoV-2 RNase P Assay Kit. The Life Sciences Testing Center COVID-19 Home Collection Kit consists of a sample label for the specific patient, anterior nasal swab, transport specimen vial (tube), instructions for use, return bag, small biohazard bag, help and contact number insert, and return labels if necessary. The clinical performance has not been established in all circulating variants but is anticipated to be reflective of the prevalent variants in circulation at the time and location of the clinical evaluation. Specimens that are collected at home will not be tested with an internal control to confirm that the specimen was properly collected. | Testing is limited to the Life Sciences Testing Center, located at 147 S Bedford St., Burlington, MA, 01803, which is certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meets the requirements to perform high-complexity tests.
ECL COVID TEST SYSTEM | Empire City Laboratories | Home collection | N/A | N/A | anterior nasal swabs | N gene, S gene, ORF1ab | N/A | The ECL COVID TEST SYSTEM is a real-time reverse transcription polymerase chain reaction (rRT-PCR) test intended for the qualitative detection of nucleic acid from SARS-CoV-2 in self- collected (unsupervised) anterior nasal swab specimens at home using the ECL COVID-19 AT HOME COLLECTION KIT. The molecular test to be used with the ECL COVID-19 AT HOME COLLECTION KIT is the cobas® SARS-CoV-2 for use on the cobas® 6800/8800 Systems- which is a real-time reverse transcription polymerase chain reaction test for the detection of SARS-CoV-2 RNA, performed at the Empire City Laboratories Inc. The clinical performance has not been established in all circulating variants but is anticipated to be reflective of the prevalent variants in circulation at the time and location of the clinical evaluation. Specimens that are collected at home will not be tested with an internal control to confirm that the specimen was properly collected. | Testing is limited to the Empire City Laboratories Inc., located at 229 49th Street, Brooklyn, NY 11220, which is certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meets requirements to perform high complexity tests.
ECL COVID TEST SYSTEM-1 | Empire City Laboratories | Home collection | N/A | N/A | anterior nasal swabs | N1 gene, ORF1ab gene | N/A | The ECL COVID TEST SYSTEM-1 is a real-time reverse transcription polymerase chain reaction (rRT-PCR) test intended for the qualitative detection of nucleic acid from SARS-CoV-2 in self- collected (unsupervised) anterior nasal swab specimens at home using the ECL COVID-19 AT HOME COLLECTION KIT. The molecular test to be used with the ECL COVID-19 AT HOME COLLECTION KIT is the EURORealTime SARS-CoV-2 which is a real-time reverse transcription polymerase chain reaction test for the detection of SARS-CoV-2 RNA, performed at the Empire City Laboratories Inc. | Testing is limited to the Empire City Laboratories Inc., located at 229 49th Street, Brooklyn, NY 11220, which is certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meets requirements to perform high complexity tests.
SalivaDirect DTC Saliva Collection Kit | Yale School of Public Health, Department of Epidemiology of Microbial Diseases | Home collection | N/A | N/A | saliva | N/A | N/A | The SalivaDirect DTC Saliva Collection Kit is a direct to consumer (DTC) product for self- collection of saliva specimens at home (which includes in a community-based setting) from individuals aged 18 years and older, including individuals without symptoms or other reasons to suspect COVID-19, sent for testing with an in vitro diagnostic (IVD) molecular test that is indicated for use with the SalivaDirect DTC Saliva Collection Kit. Negative test results are provided to the user via text, email and/or online portal. Individuals with positive, or invalid test results will be contact by a healthcare provide by phone. Individuals may request the SalivaDirect DTC Saliva Collection Kit via a patient-initiated or organization-initiated workflow. | Testing is limited to laboratories designated by Yale School of Public Health Department of Epidemiology of Microbial Diseases, that includes the Clinical Molecular Diagnostics Laboratory, Department of Pathology, Yale School of Medicine, located at 310 Cedar St., New Haven, CT 06510, certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, that meet requirements to perform high complexity tests.
SalivaDirect for use with DTC Kits | Yale School of Public Health, Department of Epidemiology of Microbial Diseases | rRT-PCR, home collection | 94.6% | 100% | saliva | N1 gene | 6000 copies/mL | SalivaDirect for use with DTC Kits is a direct to consumer product for testing saliva specimens self- collected at home (which includes in a community-based setting) using the SalivaDirect DTC Saliva Collection Kit. It can be broadly implemented as it (1) does not require saliva collection tubes containing preservatives, (2) does not require specialized equipment for nucleic acid extraction, and (3) is validated for use with products from multiple vendors. Negative test results are provided to the user via text, email and/or online portal. Individuals with positive, or invalid test results will be contacted by a healthcare provider by phone. The SalivaDirect for use with DTC Kits assay is authorized for use with the SalivaDirect DTC Saliva Collection Kit, which was authorized for use in a separate EUA (EUA210507). | Testing is limited to laboratories designated by the Yale School of Public Health, Department of Epidemiology of Microbial Diseases, that includes the Clinical Molecular Diagnostics Laboratory, Department of Pathology, Yale School of Medicine, located at 310 Cedar St., New Haven, CT 06510, that are also certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a and meet the requirements to perform high complexity tests.
Kwokman Diagnostics COVID-19 Home Collection Kit | Kwokman Diagnostics, LLC | Home collection | N/A | N/A | anterior nasal swabs | N/A | N/A | The Kwokman Diagnostics COVID-19 Home Collection Kit is intended for use to collect anterior nasal (nasal) swab specimens at home from individuals age 18 years or older (self- collected, unsupervised). After the self-collected sample is processed at the CLIA laboratory, a physician contracted with Kwokman Diagnostics reviews and approves the test results. All results are delivered via encrypted email and patients with positive and inconclusive results will receive a phone call from a health care provider licensed in the state in which the patient resides. If a specimen is received > 96 hours from the collection date, the specimen is rejected. The stability study was conducted by Gravity Diagnostics to support shipping of nasal swabs collected in 0.9% saline for up to 96 hours. | Testing is limited to laboratories designated by Kwokman Diagnostics, LLC that are certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meet requirements to perform high complexity tests and that run the specimens collected from the Kwokman Diagnostics COVID- 19 Home Collection Kit on an IVD molecular test that is indicated for use with the Kwokman Diagnostics COVID-19 Home Collection Kit when used consistent with its authorization.
Mount Sinai SARS-CoV-2 Assay | The Mount Sinai Hospital, Center for Clinical Laboratories | rRT-PCR | 100% | 100% | saliva | ORF1ab, N gene | 100 copies/mL | The Mount Sinai SARS-CoV-2 Assay is a real-time reverse transcription polymerase chain reaction (RT-PCR) test intended for the qualitative detection of nucleic acid from SARS-CoV-2 in saliva specimens that are collected at home using the Mount Sinai SARS-CoV-2 Collection Kit by any individuals age 18 years and older (self-collected), 14 years and older (self-collected under adult supervision), or 5 years and older (collected with adult assistance). The test will be used for screening students in schools and employees in business organizations. The clinical performance has not been established in all circulating variants but is anticipated to be reflective of the prevalent variants in circulation at the time and location of the clinical evaluation. | Testing is limited to the Mount Sinai COVID Laboratory, located at 281 First Avenue, 12th Floor Dazian Building, Room 1201, New York, NY 10003, which is certified under Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meets requirements to perform high complexity tests.
binx health At-Home Nasal Swab COVID-19 Sample Collection Kit | binx health Inc. | Home collection | N/A | N/A | Nasal swabs | N/A | N/A | The August update updated the authorized labeling to emphasize the need for specimnes to be shipped on the day of collection. The EUA summary was also updated to include the results of an updated specimen stability winter and summer shipping study, and extended the specimen stability to <= 120 hours from collection. The kit is for use only with a molecular test for the detection of SARS-CoV-2 RNA that is indicated for use by binx health. | Testing is limited to laboratories designated by binx health, Inc. that are CLIA certified.
SelfCheck COVID-19 TaqPath Multiplex PCR | Cleveland Clinic Robert J. Tomsich Pathology and Laboratory Medicine Institute | rRT-PCR | 100% | 100% | anterior nasal swabs | ORF1ab, N gene, and S gene | 50 GCE/mL | The SelfCheck® COVID-19 TaqPathTM Multiplex PCR is a real-time reverse transcription polymerase chain reaction (rRT-PCR) test intended for the qualitative detection of nucleic acid from SARS-CoV-2 in self-collected (unsupervised) anterior nasal swab specimens at home using the SelfCheck® COVID-19 Swabbing Kit, by individuals (18 years of age or older) suspected of COVID-19, when determined to be appropriate by a healthcare provider based on either a telemedicine visit or an in-person visit with a healthcare provider. Specimens collected using the SelfCheck® COVID-19 Swabbing Kit are transported at ambient temperature for testing at a laboratory. The molecular test to be used with the SelfCheck® COVID-19 Swabbing Kit is the TaqPathTM COVID-19 Combo Kit- which is a real-time reverse transcription polymerase chain reaction test for the detection of SARS-CoV-2 RNA, performed at the Robert J. Tomsich Pathology and Laboratory Medicine Institute at the Cleveland Clinic. | Testing is limited to the Cleveland Clinic Robert J. Tomsich Pathology and Laboratory Medicine Institute, located at 9500 Euclid Ave/LL2, Cleveland, OH 44195, which is certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meets requirements to perform high complexity tests.
Ambry COVID-19 RT-PCR Test | Ambry Genetics Laboratory | rRT-PCR | 96.8% | 100% | saliva | ORF1ab, N gene, and S gene | 100 copies/mL | The July 2021 update included new instructions for use for the Ambry Saliva Collection Kit. The March 2021 update includes the results of the FDA reference panel and removes one of the allowable extraction instruments. The Ambry COVID-19 RT-PCR Test is a real-time reverse transcription polymerase chain reaction (RT-PCR) test intended for the qualitative detection of nucleic acid from SARS-CoV-2 in saliva specimens. Saliva can be collected by a healthcare provider using the DNA Genotek OMNIgene·ORAL OM-505/OME-505 saliva collection device or at-home (for persons 18 and older) using the Ambry COVID-19 RT-PCR Test Saliva Collection Kit. Sensitivity and specificity were determined from 31 positive and 34 negative clinical samples. | Testing is limited to Ambry Genetics Laboratory located at 7 Argonaut, Aliso Viejo, CA 92656
Everlywell COVID-19 & Flu Test Home Collection Kit | Everlywell, Inc. | Home collection | N/A | N/A | anterior nasal swabs | Not stated | Not stated | The Everlywell COVID-19 & Flu Test Home Collection Kit is intended for the collection of anterior nasal swab specimens at home from individuals age 18 years and older (self- collected), 16 years and older (self-collected under adult supervision), or 2 years and older (collected with adult assistance) when suspected of respiratory viral infection consistent with COVID-19 by their healthcare provider. Specimens collected using the Everlywell COVID-19 & Flu Test Home Collection Kit are transported at ambient temperature for testing at an authorized laboratory. | Testing is limited to laboratories designated by Everlywell, Inc. that are certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meet requirements to perform high complexity tests. Testing is also limited to molecular diagnostic tests that are indicated for use with the Everlywell COVID-19 & Flu Test Home Collection Kit.
COVID-Flu Multiplex Assay | Exact Sciences Laboratories | Home collection, rRT-PCR | 100% | 100% | anterior nasal swabs | 3' terminus, N gene | 18.208 copies/mL | The COVID-Flu Multiplex Assay is a real-time RT-PCR assay intended for the simultaneous qualitative detection and differentiation of nucleic acid from SARS-CoV-2, influenza A virus, and/or influenza B virus in anterior nasal swab specimens self-collected in a healthcare setting by individuals suspected of respiratory viral infection consistent with COVID-19 by a healthcare provider. Clinical signs and symptoms of respiratory viral infection due to SARS-CoV-2 and influenza can be similar. | Testing is limited to Exact Sciences Laboratories, located at 650 Forward Drive, Madison, WI 53711, which is certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. § 263a, and meets the requirements to perform high complexity tests.
Pro-AmpRT SARS-CoV-2 Test | Pro-Lab Diagnostics | RT-LAMP | 96.6% | 100% | nasal, nasopharyngeal, and oropharyngeal swab specimens | ORF1ab | 125 genomic equivalents/swab | The May 2021 update removes the use of un-extracted samples via the direct method. The Pro-AmpRT SARS-CoV-2 Test is an RT-LAMP assay with 6 primers that amplify a region of ORF1ab. The test should be run on the Genie HT instrument with Software version v4.00rc2. Sensitivity was determined from 30 positive clinical samples and 30 negative clinical samples. | Testing is limited to Pro-Lab Diagnostics, located at 21 Cypress Blvd, Suite 1155, Round Rock, Texas 78655
Quaeris SARS-CoV-2 Assay | Harvard University Clinical Laboratory (HUCL) | Home collection, rRT-PCR | 95.7% | 100% | anterior nasal swabs | N1, RdRp | 0.0025 copies/mL | The Quaeris SARS-CoV-2 Assay is a real-time reverse transcription polymerase chain reaction (rRT-PCR) test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 virus in self-collected (unsupervised) or healthcare provider-collected anterior nasal swab specimens at home or in a healthcare setting (which includes in a community-based setting) using the COVID-19 Self-Swab Collection Kit for Harvard University Clinical Laboratory (HUCL) by individuals 18 years of age or older suspected of COVID-19. | The Quaeris SARS-CoV-2 Assay will be performed at the Harvard University Clinical Laboratory, located at B139 Northwest Laboratories, 52 Oxford Street Cambridge MA02138.
Phosphorous COVID19 RT-qPCR Test DTC | Phosphorus Diagnostics LLC | rRT-PCR | 100% | 100% | saliva | N gene (N1 and N2 regions) | 0.005 copies/mL | The Phosphorus COVID-19 RT-qPCR Test DTC is a direct to consumer product for testing of saliva specimens collected at home (which includes in a community-based setting), using the Pinpoint by Phosphorus COVID-19 Test Home Collection Kit DTC. The assay uses two primer and probe sets to detect two SARS-CoV-2 specific regions of the nucleocapsid (N) gene (the N1 and N2 targets). | The Phosphorus COVID-19 RT-qPCR Test DTC will be performed at laboratories designated by Phosphorus Diagnostics LLC that are certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a and meet requirements to perform high-complexity tests, as described in the laboratory procedures that were reviewed by the FDA under this EUA.
Pinpoint by Phosphorus COVID-19 Test Home Collection Kit DTC | Phosphorus Diagnostics LLC | Home collection, rRT-PCR | 100% | 100% | saliva | N gene (N1 and N2 regions) | 0.010 copies/mL | The Pinpoint by Phosphorus COVID-19 Test Home Collection Kit DTC is a direct to consumer (DTC) product for self-collecting (unobserved) a saliva specimen by an individual 18 years or older at home (which includes in a community-based setting), that are sent for testing with an in vitro diagnostic (IVD) molecular test that is indicated for use with the Pinpoint by Phosphorus COVID-19 Test Home Collection Kit DTC for self-collection of saliva specimens, and the IVD is indicated for testing any individuals, including individuals without symptoms or other reasons to suspect COVID-19. | Direct to consumer (DTC) home self-collected saliva specimens collected by individuals 18 years or older (unobserved) with the Pinpoint by Phosphorus COVID-19 Test Home Collection Kit DTC will be sent to high complexity laboratories that have been designated by Phosphorus Diagnostics LLC. All laboratories will be certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a to perform high complexity tests and that run specimens collected using Pinpoint by Phosphorus COVID-19 Test Home Collection Kit DTC on an in vitro diagnostic (IVD) molecular test that is indicated for use with the Pinpoint by Phosphorus COVID-19 Test Home Collection Kit DTC for self-collection of saliva specimens.
COVID-19 RT-PCR Test | Laboratory Corporation of America (LabCorp) | rRT-PCR | 100% | 100% | nasal, nasopharyngeal or oropharyngeal swabs, sputum, lower respiratory tract aspirates, bronchoalveolar lavage, and nasopharyngeal wash/aspirate or nasal aspirate; self-collected nasal swabs | Not stated | 62.5 copies/reaction | This is an update to the original EUA. The December 2020 update includes data on concordance between healthcare collection and home, self-collection. The test can now also be used for asymptomatic patients, for pooled samples (containing up to 5 pooled samples and 25 pools per matrix), updates that samples are stable for 6 days, and updated protocols to allow for the asymptomatic and pooled sampling. Also, the swab is now foam rather than cotton. The test is an rRT-qPCR test, and therefore detects viral RNA. LabCorp does not collect specimens, only running the samples. Sample extraction must be performed using the CERES Nanosciences Nanotrap Virus Capture Kit. Samples may be collected by the healthcare provider, or at home using the Pixel by LabCorp COVID-19 test home collection kit. Sensitivity and specificity were determined using contrived samples, and as such, are not presented here. However, sensitivity and specificity for asymptomatic samples were determined using true clinical samples (20 positive and 113 negative). The April 2021 update includes amendments to thermostability of swabs for use in the kit. | Testing is limited to the Center for Esoteric Testing in Burlington, NC or other laboratories designated by LabCorp that are also certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meet the requirements to perform high complexity tests.
SalivaDirect Assay | Yale School of Public Health, Department of Epidemiology of Microbial Diseases | rRT-PCR | 94.1% | 90.9% | saliva | N gene (N1 region) | 6000 copies/mL | The February 2021 update added more acceptable thermocyclers and specific cutoffs to the protocol. The December 2020 updates include the addition of a research use only designation for instrumentation used in the assay. The October 2020 updates to the EUA include an additional platform on which the test can be run, the ABI 7500 Fast Dx system. It also adds commercial resources for primer/probe sets. SalivaDirect is an RNA-extraction free, dualplex RT-qPCR method for SARS-CoV-2 detection. Saliva is first treated with proteinase K followed by a heat inactivation step, and is then directly used as input in the dualplex RT-qPCR test using validated primer and probe sets (2019-nCoV_N1 and RP) developed by the US CDC. Sensitivity and specificity was determined from 34 positive clinical samples and 33 negative clinical samples. The April 2021 update clarifies additional shipping partners for home collection samples. | Testing is limited to Yale School of Public Health, Department of Epidemiology of Microbial Diseases, that includes the Clinical Molecular Diagnostics Laboratory, Department of Pathology, Yale School of Medicine
Kaiser Permanente High Throughput SARS-CoV-2 Assay | Southern California Permanente Medical Group | Home collection, rRT-PCR | 98.6% | 96% | Saliva | N gene, S gene, ORF1ab | 525 copies/mL | The Kaiser Permanente High Throughput SARS-CoV-2 Assay is for use by prescription only, in individuals 18 and older. Individuals can request a test through their healthcare provider or through an online questionnaire. The assay uses self-collected saliva samples that are sent to the SC Permanente medical group lab. Results are given to the patient via an online portal. The test should be run on the Thermo Fisher Amplitude Solutions System with analysis on the Amplitude System COVID‑19 Interpretive Software Driver. Sensitivity was determined using 138 positive samples, and specificity using 150 negative samples. All samples were collected by a healthcare provider, and were from patients suspected of COVID-19. | Testing is limited to the Southern California Permanente Medical Group - Regional Reference Laboratory (SCPMG-RRL) located at 13000 Peyton Drive, Chino Hills, CA 91709.
Clinical Enterprise SARS-CoV-2 RT-PCR Assay | Clinical Enterprise, Inc | rRT-PCR | 100% (symptomatic), 100% (asymptomatic) | 98.4% (symptomatic), 99.1% (asymptomatic) | nasopharyngeal swabs, mid turbinate swabs, anterior nasal swabs, nasopharyngeal wash, nasal wash, oropharyngeal (OP) swabs and bronchoalveolar lavage (BAL) | N gene (N1 and N2 regions) | 250 copies/mL | The Clinical Enterprise SARS-CoV-2 RT-PCR Assay is an rRT-PCR assay for detection of SARS-CoV-2. This test can be used in individuals suspected of COVID-19 by their healthcare provider, or in asymptomatic individuals, without a prescription using anterior nasal swabs. The test should be run on the Applied Biosystems QuantStudio5 384-Well Real-Time PCR System (Software V1.5.1). The test performance was validated in retrospective and prospective (using a Boca Biolistics SARS-CoV-2 Validation panel) clinical samples, as well as clinical samples from asymptomatic individuals. Sensitivity was determined from 85 positive symptomatic samples and 22 asymptomatic positive samples. Specificity was determined using 65 symptomatic negative and 110 asymptomatic negative samples. | Testing is limited to CLIA laboratories designated by Clinical Enterprise, Inc.
Clinical Enterprise SARS-CoV-2 RT-PCR Assay DTC | Clinical Enterprise, Inc | rRT-PCR, Home collection | 100% (symptomatic), 100% (asymptomatic) | 98.4% (symptomatic), 99.1% (asymptomatic) | anterior nasal swabs | N gene (N1 and N2 regions) | 250 copies/mL | The Clinical Enterprise SARS-CoV-2 RT-PCR Assay DTC is an at-home sample collection kit partnered with an rRT-PCR assay. This test can be used in asymptomatic individuals, without a prescription, and has been validated in this population. The test should be run on the Applied Biosystems QuantStudio5 384-Well Real-Time PCR System (Software V1.5.1). The test performance was validated in retrospective and prospective (using a Boca Biolistics SARS-CoV-2 Validation panel) clinical samples, as well as clinical samples from asymptomatic individuals. Sensitivity was determined from 85 positive symptomatic samples and 22 asymptomatic positive samples. Specificity was determined using 65 symptomatic negative and 110 asymptomatic negative samples. | Direct to consumer (DTC) home collected anterior nasal (nasal) swabs collected by individuals using the EmpowerDX COVID-19 Home Collection Kit DTC must be sent to CLIA certified laboratories designated by Clinical Enterprise, Inc. and tested with the Clinical Enterprise SARS-CoV-2 RT-PCR Assay.
covidSHIELD | University of Illinois Office of the Vice President for Economic Development and Innovation | rRT-PCR | 95.8% | 98.9% | saliva | N gene, S gene, ORF1ab | 500-1000 copies/mL | The covidSHIELD assay is a direct saliva-to-RT-qPCR assay that detects three genes (ORF1ab, N-gene, and S-gene of SARS-CoV-2. The test method does not require RNA isolation, purification or extraction. Heat treatment (at 95°C for 30 minutes) and treatment with buffer TBE and detergent Tween20 disrupts viral particles and releases viral RNA. Samples must be collected by a trained observer. The April 2021 update includes that individuals may be tested who are asymptomatic if tested serially (weekly): tested at least weekly and with no more than 168 hours between tests. Sensitivity was determined using 48 positive samples and 89 negative samples from symptomatic patients. Data are unavailable for performance in asymptomatic individuals. | Testing is limited to authorized laboratories designated by the University of Illinois Office of the Vice President for Economic Development and Innovation, that includes the University of Illinois Veterinary Diagnostic Laboratory, University of Illinois Urbana Champaign School of Veterinary Medicine.
P23 Labs TaqPath SARS-CoV-2 Assay | P23 Labs, LLC | rRT-PCR | 100% | 100% | oropharyngeal swab, nasopharyngeal swab, anterior nasal and mid-turbinate nasal swabs, nasopharyngeal wash/aspirate or nasal aspirate and BAL, saliva; nasal and saliva swabs that are self-collected | N gene, S gene, ORF1ab | 10000 copies/mL | The March 2021 updates include new instructions for self-collection of swabs without telemedicine supervision. The home collection may now be performed without supervision for saliva and anterior nasal swabs, though all other types of samples must be collected by a healthcare provider. Approved home collection kits include the Everlywell COVID-19 Test Home Collection Kit and the binx Health at-home collection kit. This test is also for use with saliva specimens that are either self-collected at home or in a healthcare setting using the OMNIgene·ORAL OM-505 Collection Device. The assay uses primers and probes that were developed and validated under the EUA for the Thermo Fisher TaqPath COVID-19 Combo Kit. The P23 Labs TaqPath SARS-CoV-2 Assay is for use with the ThermoFisher Applied Biosystems QuantStudio 5 Real-Time PCR System equipped with software v1.3 (Expression Suite). Sensitivity and specificity were determined from 31 positive and 11 negative clinical samples. The October update to the EUA includes instructions for at-home saliva and nasal swab collection. | Testing is limited to P23 Labs, LLC (P23), Little Rock, AR, and CLIA laboratories designated by P23 Labs, LLC (P23).
KPMAS COVID-19 Test | Kaiser Permanente Mid-Atlantic States | Home collection | N/A | N/A | nasal swab (self collection) | N/A | N/A | The EUA was modified after the 6/13/2020 original, per Kaiser Permanente's request. The new EUA letter includes revised intended use to include testing of specimens collected from individuals without symptoms or other reasons to suspect COVID-19,. This also includes revised Conditions of Authorization to reflect more recent authorizations. The updated EUA Summary includes the data to support asymptomatic testing, and there are updated healthcare provider and patient fact sheets to reflect asymptomatic testing claims and also include language used in more recent authorizations. This KPMAS COVID-19 Test kit is for self-collection of nasal swab samples. The March 2021 update removes the requirement for telemedicine supervision of sample collection. They also performed usability and winter stability studies. | This kit is available for self collection of samples, but only through the guidance of a healthcare provider. The kit must be mailed back to Kaiser Permanente labs within 48 hours.
iC SARS-CoV-2 test | Tempus Labs, Inc. | rRT-PCR | 100% | 97% | nasopharyngeal (NP), anterior nares (AN or nasal), mid-turbinate nasal, and oropharyngeal (OP) swab specimens | N, S, and ORF1ab | 250 copies/mL | The iC SARS-CoV-2 test is an rRT-PCR based assay that is a modified version of the Applied Biosystems TaqPath COVID-19 Combo Kit. This assay differs in the collection media, extraction method, and platform used. The test should be run on the Applied Biosystem QuantStudio 7 Flex 384 instrument (software version 1.3). Sensitivity was determined from 30 positive clinical samples, and specificity from 31 negative samples (compared to the TaqPath kit results). The March 2021 update includes the addition of self-collected nasal swabs using the Tempus Nasal Sample Collection Kit in individuals 18 and older. It also included further sample stability testing. | Testing is limited to Tempus Labs, Inc.’s laboratories located at 600 W Chicago Ave, Ste 510, Chicago, IL 60654.
PacificDx Covid-19 Test | ResearchDx, Inc., DBA Pacific Diagnostics | rRT-PCR | 100% | 97% | nasopharyngeal, oropharyngeal, nasal or mid-turbinate swabs collected by a healthcare provider, or self-collected nasal swabs | N gene (N1 and N2 regions) | 1600-1900 copies/mL | This assay is a modification of the CDC diagnostic PCR-based assay. This test may be used in those 18 and over with the RapidRona Self Collection Kit. The test should be run on the Applied Biosystems QuantStudio Dx or 7900HT platform. Sensitivity was determined using 89 positive clinical samples, and specificity using 67 negative samples. | Testing is limited to Pacific Diagnostics located at 5 Mason Drive, Irvine, CA 92618
Infinity BiologiX TaqPath SARS-CoV-2 Assay | Infinity BiologiX, LLC | rRT-PCR | 100% | 100% | oropharyngeal swab, nasopharyngeal swab, anterior nasal swab, mid-turbinate nasal swab, and bronchoalveolar lavage (BAL) fluid; saliva from at-home collection | N gene, S gene, ORF1ab | 200 copies/mL | The March 2021 update includes a new authorized distributer, Wellness 4 Humanity. November updates to the EUA include the addition of authorized distributors other than Infinity Biologix.This test was previously known as the “Rutgers Clinical Genomics Laboratory TaqPath SARSCoV-2 Assay." This is an update to the existing EUA, primarily in the transfer of EUA from Rutgers to Infinity BiologiX LLC and update the device name to Infinity BiologiX TaqPath SARS-CoV-2 Assay Saliva specimens may also be collected for use with your product using the Spectrum Solutions LLC SDNA-1000 Saliva Collection Device to self-collect salvia specimens at home or in a healthcare setting. This assay is for use with the ThermoFisher Applied Biosystems QuantStudio 5 Real-Time PCR System equipped with software v1.3, or the Applied Biosystems ViiA7 Real-Time PCR System with the Applied Biosystems QuantStudio 5 software v1.3 for data analysis, and Perkin Elmer Chemagic 360 extraction instrument (software v6.3.0.3). The device is also under the names: Vault Health COVID-19 Test, USGC Saliva Test Kit, hims/hers Saliva Test Kit, BioIQ COVID-19 Saliva Test, Concentric by Ginkgo, 1health.io Zero Contact, Base10 Saliva Test Kit, Purlab. Saliva Test Kit, Reliant Saliva Test Kit, E7 Health Saliva Test Kit, miinehealth Saliva Test Kit, Maverick Saliva Test Kit, SDTL Saliva Test Kit, and AZOVA At-Home COVID-19 Test Collection Kit. | Testing is limited to Infinity BiologiX LLC, Nelson Biological Sciences, C-Wing, 604 Allison Road, Piscataway, NJ 08854
Gravity Diagnostics SARS-CoV-2 RT-PCR for use with DTC kits | Gravity Diagnostics, LLC | Home collection, rRT-PCR | 96.4% (symptomatic), 100% (asymptomatic) | 100% (symptomatic), 100% (asymptomatic) | nasal swab | N, ORF1ab, S genes | 500-1000 copies/mL | This test is for use with the 1) Everlywell COVID-19 Test Home Collection Kit DTC (18 and older) or 2) Kroger Health COVID-19 Test Home Collection Kit (16 and older) and 3) the GetMyDNA COVID-19 Home Collection Kit. The test should be run on the Applied Biosystems QuantStudio7 Flex instrument with software version 1.3, or a QuantStudio12 Flex instrument with softwareversion 1.2.2. The test was validated in both symptomatic and asymptomatic individuals using the Everlywell kit. There were 55 positive and 115 negative clinical symptomatic samples, and 20 positive and 100 negative clinical samples from asymptomatic individuals. | Testing is limited to Gravity Diagnostics, LLC, located at 632 Russell Street, Covington, KY 41011 and 812 Russell Street, Covington, KY 41011
GetMyDNA COVID-19 Test Home Collection Kit | GetMyDNA | Home collection | N/A | N/A | anterior nasal swabs | N/A | N/A | The GetMyDNA COVID-19 Home Collection kit is for the unsupervised collection (in people 18 years and older) of anterior nasal swabs for molecular diagnostic testing. This collection kit is authorized for use only with molecular diagnostics designated by GetMyDNA. One of these designated diagnostic tests is the Gravity Diagnostics SARS-CoV-2 RT-PCR test. This kit can be used in people who are asymptomatic. Test results are given to the user via an online portal, and those who test positive (or inconclusive) will be contacted by a healthcare provider. Sample stability was tested for 48-96 hours, and they performed usability studies. | Testing is limited to high complexity CLIA laboratories that have been designated by GetMyDNA.
Phosphorus COVID-19 RT-qPCR Test | Phosphorous Diagnostics, LLC | rRT-PCR | 97.1% | 98.2% | oropharyngeal (throat) swabs, nasopharyngeal swabs, anterior nasal and mid-turbinate nasal swabs, nasopharyngeal washes/aspirates or nasal aspirates as well as bronchoalveolar lavage (BAL), and self-collected saliva samples | N gene (N1 and N2 regions) | 5000 copies/mL | The March 2021 updates include language and title changes for the test collection methods (either by HCP or self collected). The December 2020 updates include additional types of samples acceptable, an additional authorized laboratory for testing, and updated language. Saliva specimens can either be self-collected at home or in a healthcare setting, both using the Oragene Dx OGD-510 kit. The Phosphorus COVID-19 RT-qPCR Test is for use with the CFX384 Touch Real-Time PCR Detection System with Bio-Rad CFX Manager software version 3.1. | Testing is limited to Phosphorus Diagnostics LLC located in Secaucus, NJ and Wheeler Labs located at 655 Research Parkway, Suite 383, Oklahoma City, OK 73104
CRSP SARS-CoV-2 Real-time Reverse Transcriptase (RT)-PCR Diagnostic Assay | Clinical Research Sequencing Platform (CRSP), LLC at the Broad Institute of MIT and Harvard | rRT-PCR | 98.3% | 100% | anterior nasal swab | N gene (N1 and N2 regions) | 1600 copies/mL | The CRSP SARS-CoV-2 Real-time Reverse Transcriptase (RT)-PCR Diagnostic Assay (Version 3) is also authorized for the qualitative detection of nucleic acid from SARS-CoV-2 in dry anterior nasal swab specimens self-collected unsupervised using the CRSP Self-Swab kit, by individuals (18 years of age and older) suspected of COVID-19 by their healthcare provider and when determined to be appropriate by the healthcare provider. The test is also authorized for use with the Color COVID-19 Self-Swab collection Kit for individuals (18 years of age and older) to self-collect anterior nasal swab specimens unsupervised at home. This test uses primer/probe sets developed by the CDC that target two viral gene targets in the Nucleocapsid gene of SARSCoV-2, N1 and N2, and an internal control gene, RNase P (RP) in a multiplexed reaction. | Testing is limited to the Clinical Research Sequencing Platform (CRSP), LLC at the Broad Institute of MIT and Harvard located at 320 Charles Street, Cambridge, MA 02141.
SARS-CoV-2 RT-PCR test | Viracor Eurofins Clinical Diagnostics | rRT-PCR | 100% | 99.2% | nasopharyngeal swab, nasal swab, nasopharyngeal wash, nasal wash, oropharyngeal swab, BAL, and at-home nasal swabs | Nucleoprotein (two targets - N1 and N2) | 73 copies/mL | The Feb. 26, 2021 reauthorization modifies the test as a "direct to consumer" product. The updates to the EUA in October 2020 include allowing pooling of specific clinical specimens and including testing of nasal swab specimens collected with the EmpowerDX At-Home COVID-19 PCR Test Kit. The Viracor SARS-CoV-2 assay is to be used with the bioMérieux NUCLISENS® easyMAG® and bioMérieux EMAG® for automated nucleic acid extraction and the Applied Biosystems™ 7500 Real-Time PCR Systems (SDS v1.5.1) for nucleic acid amplification and detection. The test is run as a multiplex reaction with SARS-CoV-2 N1, SARS-CoV-2 N2 and MS2 internal control assays combined in a single tube. Since both SARS-CoV-2 N1 and SARS-CoV-2 N2 assays use probes with the same fluorophore (FAM), a single SARS-CoV-2 CT value is generated and interpreted for each rRT-PCR reaction. | Healthcare providers can send clinical specimens for testing to Viracor Eurofins Clinical Diagnostics
DxTerity SARS-CoV-2 RT PCR CE Test | DxTerity Diagnostics, Inc | rRT-PCR | 97.3% | 90% | saliva | N gene, E gene, ORF1ab | 50 copies/mL | The DxTerity SARS-CoV-2 RT-PCR CE Test is an end point reverse transcription polymerase chain reaction (RT-PCR) test followed by detection with capillary electrophoresis (CE). Saliva specimens must be self-collected using the DxTerity COVID-19 Test Collection Kit. The company performed sample stability studies at high and low temperatures for up to 56 hours. Recipients fill out an online patient questionnaire and the kit will be shipped to any individuals for COVID-19 testing as determined to be appropriate by the health care provider, including from individuals without symptoms of COVID-19. Sensitivity and specificity was determined from 37 positive clinical samples and 40 negative clinical samples. Performance has not been validated using all circulating variants. | Testing is limited to DxTerity Diagnostics, Inc. located at 19500 S. Rancho Way, Suite 116, Rancho Dominguez, CA 90220
Kroger Health COVID-19 Test Home Collection Kit | The Kroger Co. | Home collection | N/A | N/A | nasal swab |  |  | The February 2021 update includes the direct to consumer option for purchasing the test. Individuals do not have to have symptoms or be suspected of COVID-19. It can be used in individuals 16 and older. The at-home collection kit uses a saline solution to stabilize the nasal swab sample for processing. The collection will be performed by the patient, with or without guidance over video from a healthcare provider. They tested for sample stability over time, at various temperatures, and also tested for compliance to the self-collection protocol. They have stated that even after release, they will continue to monitor compliance and performance of the test kit. They will be partnering with Gravity Diagnostics to run the tests, which received EUA on June 1, 2020. | Testing is limited to CLIA laboratories designated by The Kroger Co. and meet requirements to perform high complexity tests and that run the specimens collected from the Kroger Health COVID-19 Test Home Collection Kit on an in vitro diagnostic (IVD) molecular test that is indicated for use with the Kroger Health COVID-19 Test Home Collection Kit for self-collection of nasal swab specimens as described above | Kroger has stated in a press release that they will process 60,000 tests per week by the end of July. The test will first be available to frontline associates at Kroger stores.
Assurance SARS-CoV-2 Panel DTC | Assurance Scientific Laboratories | Home collection, rRT-PCR | 95.2% (symptomatic), 95.4% (asymptomatic) | 97.7% (symptomatic), 97.5% (asymptomatic) | nasal swab | N gene (N1) | 9000 copies/mL | The Assurance SARS-CoV-2 Panel DTC is a direct to consumer test to be used with 1) the Simplicity COVID-19 Home Collection Kit or (2) the Everlywell COVID-19 Test Home Collection Kit DTC by individuals 18 and older. This test can be performed unsupervised, and in asymptomatic individuals. Individuals ordering online are recommended to also complete a questionnaire but it is not required. Test results will be communicated via an online portal, and a healthcare provider will contact the individual if results are positive or indeterminate. The panel was tested in both symptomatic and asymptomatic individuals using the Everlywell kit. There were 62 positive and 217 negative clinical symptomatic samples, and 43 positive and 199 negative clinical samples from asymptomatic individuals. The NIH reference panel gave an LOD of 5400 copies/mL. | Testing is limited to CLIA laboratories designated by Assurance Scientific Laboratories.
Everlywell COVID-19 Test Home Collection Kit DTC | Everlywell, Inc. | Home collection | Not stated | Not stated | nasal swab | Not stated | Not stated | This collection kit is a direct to consumer test. This is for at-home use, unsupervised collection of nasal swabs by individuals 18 and older. This can be used by asymptomatic individuals. Everlywell performed self-collection validation studies, as well as stability studies in high or low temperatures. This collection kit can only be sent to labs designated by Everlywell, Inc. for processing. Test results will be communicated via an online portal, and a healthcare provider will contact the individual if results are positive or indeterminate. | Testing is limited to CLIA laboratories designated by Everlywell, Inc.
QDX SARS-CoV-2 Assay | QDx Pathology Services | Home collection | Not stated | Not stated | nasal swab | Not stated | Not stated | The QDX SARS-CoV-2 Assay primarily focuses on at-home self-collection of nasal swabs. These are then sent to the QDx laboratory for testing. Testing is limited to individuals 18 and older. The February update includes winter and summer temperature stability studies, which indicated that samples would be stable at low (-10C) and high (40C) temperatures for several hours. | Testing is limited to the QDx Pathology Services laboratory located at 46 Jackson Drive, Cranford, NJ 07016.
Cleveland Clinic SARS-CoV-2 Assay | Cleveland Clinic Robert J. Tomsich Pathology and Laboratory Medicine Institute | rRT-PCR | 97% | 100% | nasopharyngeal (NP), oropharyngeal (OP), and nasal swabs and bronchoalveolar lavage (BAL), tracheal aspirate (TASP), and sputum | E gene and RDRp | 10,000 copies/mL | The Cleveland Clinic SARS-CoV-2 Assay is a multiplexed rRT-PCR assay. The assay can also be coupled with the SelfCheck COVID-19 Swabbing Kit in individuals 18 years of age and older. The January 2021 update the allowance for testing of specimens collected at home (by prescription only). This test should be run on the Roche LightCycler 480 and/or cobas z 480 platform. Sensitivity and specificity were determined from 30 positive and 30 negative nasopharyngeal clinical samples. They also tested sensitivity on a limited number of other types of specimens, which are listed in the IFU (all had 100% sensitivity and specificity). 20 sputum (10 positive and 10 negative), 22 BAL (11 positive and 11 negative), and 31 tracheal aspirates (TASP) (13 positive and 18 negative) samples were also tested. Sputum, BAL, and TASP samples all had 100% sensitivity and specificity. | Testing is limited to the Robert J. Tomsich Pathology and Laboratory Medicine Institute located at 9500 Euclid Ave./LL2, Cleveland, OH 44195
Curative-Korva SARS-Cov-2 Assay | KorvaLabs Inc. | rRT-PCR | 100% | 100% | oropharyngeal (throat) swab, nasopharyngeal swab, nasal swab, and oral fluid | N gene (N1 and N2) | 200 copies/mL | This test should only be performed within 14 days of symptom onset, collection must be performed by a healthcare provider, and only on symptomatic patients. The test should be run on the BioRad CFX 96 Touch, BioRad CFX Connect Real-Time PCR systems and Roche LightCycler 480 II Real-Time PCR systems. Sensitivity was determined from 5 positive samples, and specificity from 5 samples. In January, the FDA released a warning against false negative results with this assay. | CLIA approved laboratories are authorized to use the kit. It is approved for in vitro diagnostic use under FDA Emergency Use Authorization only. It is available for purchase by healthcare/research professionals. | According to their website (date not included): "Our laboratory is set up for high throughput and can currently accommodate up to 10,000 samples per day but rapidly gearing up for more to service as many cities, counties and states as possible." | https://www.fda.gov/medical-devices/safety-communications/risk-false-results-curative-sars-cov-2-test-covid-19-fda-safety-communication
Quest Diagnostics RC COVID-19 +Flu RT-PCR | Quest Diagnostics Infectious Disease, Inc | rRT-PCR | 96.4% (SARS-CoV-2), 100% (influenza A), 100% (influenza B) | 98% (SARS-CoV-2), 99.6% (influenza A), 99.7% (influenza B) | nasal swabs (self-collected or by healthcare professional) | Not stated | ~730 copies/mL for SARS-CoV-2 | The Quest Diagnostics RC COVID-19 +Flu RT-PCR is intended for the simultaneous qualitative detection and differentiation of nucleic acid from SARS-CoV-2, influenza A virus and/or influenza B virus in nasal swab specimens self-collected at home, using the Quest Diagnostics Self-Collection Kit for COVID-19 +Flu. Self-collected nasal swab specimens will be tested using the Quest Diagnostics RC COVID-19 +Flu RT-PCR which is performed using the FDA EUA-authorized Roche cobas SARS-CoV-2 & Influenza A/B molecular test, which is an automated RT-PCR based platform. The test report will then be electronically delivered to both the ordering healthcare provider and the participant. | Testing is limited to laboratories designated by Quest Diagnostics that are certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, and meet the requirements to perform high complexity tests.
SARS-CoV-2 RT-PCR Assay | Stanford Health Care Clinical Virology Laboratory | rRT-PCR | 100% | 100% | nasopharyngeal, oropharyngeal, nasal, and mid turbinate nasal swabs, BAL | E gene, human RP | 1000 copies/mL | The 2019 novel Coronavirus (SARS-CoV-2) Real-Time, RT-PCR test is to be used with the Qiagen Rotor-Gene Q thermocycler using Rotor-Gene Q Series Software 2.3.1 (Build 49). The following extraction platforms can be used: Qiagen EZ1 XL and QIAsymphony SP/AS. Revision included minimal changes to EUA summary. | Healthcare providers can send clinical specimens for testing to Stanford Health Care Clinical Virology Laboratory
Gravity Diagnostics SARS-CoV-2 RT-PCR Assay | Gravity Diagnostics, LLC | RT-PCR | 96.4% | 100% | nasal, nasopharyngeal (NP), and oropharyngeal (OP) swab specimens (for any patient) saliva (for symptomatic patients only) | N, ORF1ab, S genes | 500-1,000 GE/mL (depending on instrument used) | Specimens can be collected by a healthcare provider or at-home using designated EUA self-collection kits. The Gravity Diagnostics SARS-CoV-2 RT-PCR assay is a real-time reverse transcription polymerase chain reaction (rRT -PCR) test. The test uses three primer and probe sets to detect one region in the N gene, one region in the ORF1ab, and one region in the S gene. The Gravity Diagnostics SARS-CoV-2 assay is to be used with the Thermo Fisher KingFisher Flex with software version 2.5 and the Hamilton Microlab STAR with software version 1.01 automated nucleic acid isolation instruments. | Healthcare providers can send clinical specimens for testing to Gravity Diagnostics, LLC, located at 632 Russell Street, Covington, KY 41011 and 812 Russell Street, Covington, KY 41011
RapidRona Self Collection Kit | RapidRona, Inc. | Home collection |  |  | nasal swab |  |  | The RapidRona Self-Collection Kit is intended for use by individuals for self-collection of nasal swab specimens at home, when determined by a healthcare provider to be appropriate based on the results of an online COVID-19 questionnaire. Specimens are to be mailed to the laboratory for testing using the prelabeled return envelope. Each RapidRona Self-Collection Kit is intended to be picked up on the day of specimen collection for return shipment overnight at ambient conditions for next day delivery. | The self-collection kit can be sent to High Complexity CLIA laboratories that have been designated by RapidRona Inc. to conduct the testing.
CRSP SARS-CoV-2 Real-time Reverse Transcriptase (RT)-PCR Diagnostic Assay | Clinical Research Sequencing Platform (CRSP), LLC at the Broad Institute of MIT and Harvard | rRT-PCR | 100% (both versions) | 100% (both versions) | nasopharyngeal, oropharyngeal, nasal, mid-turbinate swabs, nasopharyngeal wash/aspirate or nasal aspirate specimens, and bronchoalveolar lavage specimens; self collected nasal swabs | N gene (N1 and N2 regions) | Version 1-4000 copies/mL; Version 2- 1600 copies/mL | The October update includes revisions to clarify the verion of test to use, as well as inclusion of at-home self-collected dry nasal swab samples (using Kit version 2). This test has two versions, version 1 which tests for N1 and N2, and the human RNAseP gene, and version 2 tests for N2 and RNAseP. The primer and probe sequences are the same as those published by the CDC. Version 1 should be run on the Applied Biosystems Viia7 thermocycler or the Applied Biosystems QuantStudio 7 Flex (each with QuantStudio software version 1). Version 2 can only be run on the AB QuantStudio 7 Flex.Version 1 sensitivity was determined from 33 positive clinical samples, and specificity from 40 negative samples. Of these, there were thirteen positive oropharyngeal (OP) swabs, 10 positive nasopharyngeal (NP) swabs, 10 positive NP/OP swabs, and 40 negative NP swabs. Version 2 sensitivity was determined from 30 positive nasopharyngeal samples and specificity from 28 negative nasopharyngeal samples. | Testing is limited to the Clinical Research Sequencing Platform (CRSP), LLC at the Broad Institute of MIT and Harvard located at 320 Charles Street, Cambridge, MA 02141
WREN Laboratories COVID-19 PCR Test | Wren Laboratories, LLC | rRT-PCR | 100% | 95% | nasopharyngeal, oropharyngeal (throat), anterior nasal, and midturbinate nasal swabs, as well as nasopharyngeal washes/aspirates or nasal aspirates, and bronchoalveolar lavage (BAL) specimens, and saliva specimens (self collected) | N1 of SARS-CoV-2, and N3 of Sarbecovirus | 10,000 copies/mL | The October update includes at-home saliva sample collection. Saliva specimens may be self-collected at home or in a healthcare setting using the WREN Laboratories Saliva Collection Kit. The Wren Laboratories COVID-19 PCR Test is an rRT-PCR based assay. The test should be run on Applied Biosystems QuantStudio 7-Flex Real-Time PCR Instrument with QuantStudio Real-Time PCR software v1.3. Sensitivity and specificity were determined from 60 positive and 60 negative clinical samples. | Testing is limited to Wren Laboratories located at 688 East Main Street, Branford, CT, 06405
Genetrack SARS-CoV-2 Molecular Assay | Genetrack Biolabs, Inc | rRT-PCR/at home collection | 100% | 100% | anterior nasal swab specimens self-collected at home | N gene (N1 and N2 regions) | 1000 copies/mL | The October update includes winter shipping stability data. The Genetrack SARS-CoV-2 Molecular Assay should be used with an at-home collection kit, the Vo' COVID-19 Test Home Collection Kit. The kit must be returned within 48 hours of sample collection. The test should be run on the pplied Biosystems QuantStudio 6 Pro Real-Time PCR System equipped with software v2.4.1. The primer and probe sets are identical to the CDC primer/probe published sets. Sensitivity and specificity were determined using 30 positive and 30 negative samples. | Testing is limited toGenetrack Biolabs Inc., located at 100 - 2806 Kingsway Vancouver BC V5R 5T5 Canada
EmpowerDX At-Home COVID-19 PCR Test Kit | Clinical Enterprise, Inc | Home collection | N/A | N/A | nasal swabs | N/A | N/A | The EmpowerDX At-Home COVID-19 PCR Test Kit is an at-home collection kit for nasal swab samples. The samples must be mailed via FedEx the day of collection. | Testing is limited to laboratories designated by Clinical Enterprise, Inc. that are CLIA certified
UCLA SwabSeq COVID-19 Diagnostic Platform | University of California, Los Angeles (UCLA) | NGS | 100% | 100% | nasopharyngeal swabs, oropharyngeal (throat) swabs, mid-turbinate nasal swabs and anterior nasal swabs) | S2 region | 250 genomic copies/mL | The UCLA SwabSeq COVID-19 Diagnostic Platform uses next generation sequencing technologies to detect SARS-CoV-2 infections. The assay has 3 primer targets: the S2 region of spike, and internal S2 control that differs from the target by 6 nt, and a human RPP30 internal control. The test should be run with the Illumina MiSeq sequencer, followed by RTA and Illumina bcl2fastq. Sensitivity was determined using 31 positive clinical samples, and specificity using 35 clinical samples. | Testing is limited to UCLA Clinical Microbiology Laboratory located at Brentwood Annex, 11633 San Vicente Blvd. Los Angeles, CA 90049 and UCLA SwabSeq COVID-19 Testing Laboratory located at 650 Charles E. Young Dr. South, CHS Building, South Tower, 87-300 Los Angeles CA, 90095,
Cuur Diagnostics SARS-CoV-2 Molecular Assay | Cuur Diagnostics | rRT-PCR | 96% | 100% | nasopharyngeal, oropharyngeal, mid-turbinate, and nasal swab | N gene, S gene, ORF1ab | 25,000 copies/mL | The update to the EUA on 10/2/2020 changed the authorized labs to perform the test, granting that Cuur Diagnostics could authorize a laboratory to run its assay. The Cuur Diagnostics SARS-CoV-2 Molecular Assay is a modified version of the Thermo Fisher TaqPath COVID-19 Combo Kit, in that it uses a 384 well plate (higher throughput) and uses a different platform. It should be run on the Applied Biosystems QuantStudio 12K Flex platform, and should be analyzed using the Applied Biosystems QuantStudio 12K Flex Gene Expression software, v1.4. Sensitivity was determined from 50 positive clinical samples, and specificity from 50 negative samples. | Testing is limited to laboratories designated by Cuur Diagnostics, located at 8876 Spanish Ridge Avenue, Suite 203, Las Vegas, Nevada
iC SARS-CoV-2 test | Tempus Labs, Inc. | rRT-PCR | 100% | 97% | nasopharyngeal (NP), anterior nares (AN or nasal), mid-turbinate nasal, and oropharyngeal (OP) swab specimens | N, S, and ORF1ab | 250 copies/mL | The iC SARS-CoV-2 test is an rRT-PCR based assay that is a modified version of the Applied Biosystems TaqPath COVID-19 Combo Kit. This assay differs in the collection media, extraction method, and platform used. The test should be run on the Applied Biosystem QuantStudio 7 Flex 384 instrument (software version 1.3). Sensitivity was determined from 30 positive clinical samples, and specificity from 31 negative samples (compared to the TaqPath kit results). | Testing is limited to Tempus Labs, Inc.’s laboratories located at 600 W Chicago Ave, Ste 510, Chicago, IL 60654 and 3155 Northwoods Place, Peachtree Corners, GA 30071
Aeon Global Health SARS-CoV-2 Assay | Aeon Global Health | rRT-PCR | 100% | 97% | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swab specimens, nasopharyngeal wash/aspirate or nasal aspirate specimens, and bronchoalveolar lavage | N, S, and ORF1ab | 250 copies/mL | The Aeon Global Health SARS-CoV-2 Assay was developed using the Thermo Fisher Scientific TaqPath COVID-19 Combo Kit. Sensitivity was determined from 35 positive clinical samples, and specificity from 35 negative samples. | Testing is limited to Aeon Global Health laboratory located at 2225 Centennial Drive, Gainesville, GA 30504
Alimetrix SARS-CoV-2 RT-PCR Assay | Alimetrix, Inc. | rRT-PCR | 100% (automated extraction), 96.7% (manual) | 93.3% (automated extraction), 93% (manual) | nasopharyngeal, oropharyngeal, anterior nasal and mid-turbinate nasal swab specimens, as well as nasopharyngeal washes/aspirates or nasal aspirates and bronchoalveolar lavage | ORF1ab, N gene (N1 and N2 regions) | 250-500 copies/mL | The Alimetrix SARS-CoV-2 RT-PCR Assay uses rRT-PCR and microarrays to detect SARS-CoV-2 infection. The microarray protocol is performed by a robot. The LoD depends on the extraction kit used. Sensitivity was determined from 30 positive clinical samples, and specificity from 30 negative samples. | Testing is limited to the Alimetrix, Inc. located at 800 Hudson Way, Suite 2200, Huntsville, AL 35806
CentoSure SARS-CoV-2 RT-PCR Assay | Centogene US, LLC | rRT-PCR | 100% | 100% | dry oropharyngeal swab specimens from patients without symptoms | N gene (N1 and N2 regions) | 2000 copies/mL | The CentoSure SARS-CoV-2 RT-PCR Assay is intended for use in asymptomatic patients. The test should be run on the Roche LightCycler 480 II platform. Sensitivity was determined using 25 positive clinical samples, and specificity from 100 negative samples. | Testing is limited to the laboratories designated by CENTOGENE US. LLC located at 99 Erie Street, Cambridge MA 02139
Akron Children’s Hospital SARS-Cov-2 Assay | Akron Children's Hospital | rRT-PCR | 97.1% | 98% | nasopharyngeal swabs, oropharyngeal (throat) swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal aspirates, nasal washes and bronchoalveolar lavage | Spike and E gene | 250 copies/mL | The Akron Children’s Hospital SARS-CoV-2 Assay is an rRT-PCR based diagnostic. The test should be run on the QIAGEN RotorGene Q MDx platform. The test is equivalent to the Altona Diagnostics RealStar SARS-CoV-2 RT-PCR Kit 1.0. They also tested stability of the sample in various alternative transport medias. Sensitivity was determined using a total of 30 samples, and specificity from 61 negative samples. These samples were from 3 different phases of testing, so sensitivity and specificity reported here are the average of all trials. | Testing is limited to the Akron Children’s Hospital located at One Perkins Square, Akron, OH 44308
AvellinoCoV2 test | Avellino Lab USA, Inc. | rRT-PCR | 100% | 100% | nasopharyngeal and oropharyngeal swabs | N gene (2 targets) | 55,000 copies/mL | The test should be run on Applied Biosystems 7500 Fast Real-Time PCR System with software version 2.3. The LoD according to the NIH reference panel is 18,000 copies/mL. Sensitivity and specificity were determined using 30 positive and 30 negative clinical samples. | Testing is limited to Avellino Lab USA, Inc.,
Assurance SARS-CoV-2 Panel | Assurance Scientific Laboratories | rRT-PCR | 100% | 100% | nasal, nasopharyngeal or oropharyngeal swabs (collected by healthcare provider) or nasal swabs (self collected) | N gene (N1) | 9000 copies/mL | The August update to the EUA included the authorization of laboratories to include laboratories designated by Assurance Scientific Laboratories that are also certified under CLIA, and also including the Assurance Home Collection Kit in the instructions for use. Also, the N2 primers and probes were removed from the kit.This test can be used in conjuction with the Everlywell at home testing kit, and should be run using the BioRad CFX96 or CFX384 platforms. The kit includes internal control primer and probe sets, and the primer and probe sequences are from the published CDC set. They determined sensitivity and specificity using contrived samples (data not presented here), as well as 5 positive and 5 negative clinical samples. The LoD from the NIH reference panel is 5400 copies/mL. | Testing is limited to laboratories designated by Assurance Scientific Laboratories that are also certified under the Clinical Laboratory Improvement Amendments
Gravity Diagnostics COVID-19 Assay | Gravity Diagnostics, LLC | rRT-PCR | 96.8% | 100% | nasal, nasopharyngeal (NP) and oropharyngeal (OP) swab and Bronchoalveolar Lavage (BAL), and home-collection of nasal swabs | N gene (N1 and N2 regions) | 2400 copies/mL | The Gravity Diagnostics COVID-19 Assay can be used with home- or healthcare setting-collected samples. The test uses the same primer and probe sets as published by the CDC. Sensitivity and specificity presented here are from true clinical specimens (31 positive and 65 negative), though contrived samples were used in a separate study. The LoD as measured by the NIH reference panel was 18,000 copies/mL. | Testing is limited to Gravity Diagnostics, LLC
Acupath COVID-19 Real-Time (RT-PCR) Assay | Acupath Laboratories, Inc. | rRT-PCR | 100% | 100% | nasopharyngeal swab, nasopharyngeal aspirate, and bronchoalveolar lavage | ORF1ab, N gene, Spike gene | 25,000 copies/mL | The Acupath COVID-19 Real-Time (RT-PCR) Assay uses the TaqPath RT-PCR COVID-19 Kit and multiplexed on the QuantStudio 12K Flex instrument. Sensitivity was determined using 30 positive clinical samples, and specificity using 30 negative samples. | Testing is limited to the Acupath Laboratories
Boston Heart COVID-19 RT-PCR Test | Boston Heart Diagnostics | rRT-PCR | 96.7% | 96.7% | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swab specimens, nasopharyngeal aspirate and bronchoalveolar lavage | N, S, and ORF1ab | 125 copies/mL | The update on 9/21 includes a protocol for using the Agilent Bravo automated liquid handler. This rRT-PCR test assays for three different viral genes. The test uses a human internal control gene of MS2. Sensitivity was determined from 30 positive clinical samples, and specificity from 30 negative samples. | Testing is limited to the Boston Heart Diagnostics laboratory located at 200 Crossing Blvd., Framingham, MA 01702
Lilly SARS-CoV-2 Assay | Eli Lilly and Company | rRT-PCR | 100% | 93% | nasopharyngeal swabs, oropharyngeal (throat) swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal aspirates, nasal washes and bronchoalveolar lavage | N gene (N1 and N2 regions) | 1000 copies/mL | Updates to the original EUA include a post-EUA clinical performance trial, which included 15 more positive and 15 more negative clinical samples (the sensitivity and specificity were both 100%). The test uses the primer and probes sets as published by the CDC. It should be run on the Applied Biosystems platform. The lab also tested the protocol on several other instrumentation options. While the lab tested performance against contrived samples, sensitivity and specificity listed here were determined from 14 positive and 15 negative true clinical samples. | Testing is limited to the Eli Lilly and Company Clinical Diagnostics Laboratory located at Lilly Corporate Center 98A/3C, Indianapolis, IN 46285
Ethos Laboratories SARS-CoV-2 MALDI-TOF Assay | Ethos Laboratories | Home collection, rRT-PCR and MALDI-TOF | 98.1% | 96.3% | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swab specimens) and bronchoalveolar lavage specimens; nasal specimen (self collect) | N gene( N1, N2, N3,) ORF1, and ORF1ab | 1 TCID50/mL | The update on 9/21 includes the instructions for use for at-home collection. The at-home collection kit must be prescribed. Ethos Laboratories SARS-CoV-2 MALDI-TOF Assay uses the Agena SARS-CoV-2 Panel Set for amplification of SARS-CoV-2 RNA. Short amplicon primer treatments rid the sample of any free nucleotides, and then terminator nucleotides are added to amplified DNA. These amplicons then have different masses, which are detected by MALDI-TOF. MassARRAY System can then be used to analyze the mass and determine the corresponding genetic sequence. This allows for many targets to be detected. Sensitivity was determined from 53 positive clinical samples, and specificity from 54 negative samples. | Testing is limited to the Ethos Laboratory located at 29 E. 6th St., Newport, KY 41071
SARS-CoV-2 (N gene detection) Test | Exact Sciences Laboratories | rRT-PCR | 95% | 100% | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swab specimens | N gene (two targets) | 2600 copies/mL | The EUA was recently updated to include the Exact Sciences Home Collection Kit for Nasal swab Collection, and account for this in the instructions for use. The updated IFU also has some changes in the procedures. The test uses specific primer and probe sets developed by CDC to detect two regions (nCOV_N1 and nCov_N2) in the SARS-CoV-2 nucleocapsid (N) gene and one primer and probe set to detect human RNase P (RP) in a clinical sample. This test uses the ABI 7500 Fast Dx (Applied Biosystems) instrument with Sequence Detection Software (SDS) version 1.4.1. Sensitivity and specificity were determined from 40 positive and 38 negative clinical samples. | Healthcare providers can send clinical specimens for testing to the Exact Sciences Laboratories laboratory
Verily COVID-19 RT-PCR Test | Verily Life Sciences | rRT-PCR, pooling | 100% | 100% | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swab specimens | ORF1ab, N gene, Spike gene | 60 copies/mL | The Verily COVID-19 RT-PCR Test is a rRT-PCR based test that allows for pooling of up to 12 individual samples per run. The test uses a modified version of the ThermoFisher Scientific TaqPath COVID-19 Combo Kit, which previously received EUA. The test uses a 2D matrix pooling strategy within a 96 well plate. Sensitivity was determined from 35 positive clinical samples, and specificity from 30 negative clinical samples. | Testing is limited to Verily Life Sciences laboratory, located at 249 E Grand Avenue, South San Francisco, CA 94080
MiraDx SARS-CoV-2 RT-PCR assay | MiraDx | rRT-PCR | 96.9% | 100% | nasal, mid-turbinate, nasopharyngeal, oropharyngeal swab specimens and nasopharyngeal wash/aspirate or nasal aspirate specimens | N gene | 4000 copies/mL | The MiraDx SARS-CoV-2 RT-PCR detects SARS-CoV-2 infection through a single primer target, the nucleocapsid gene. The test should be run on the Applied Biosystems QuantStudio 6 Flex PCR Instrument. Sensitivity was determined from 32 positive clinical samples, and specificity from 37 negative samples. | Testing is limited to MiraDx located at 11600 Wilshire Blvd., Suite 410, Los Angeles, CA 90025
BayCare SARS-CoV-2 RT PCR Assay | BayCare Laboratories, LLC | rRT-PCR |  |  | nasopharyngeal swab specimens and pooled samples (with up to 2 patient swabs per pool) | ORF1, E gene | 0.009 TCID50/mL * did not state the TCID50 concentration | This test uses the cobas SARS-CoV-2 kit, which previously received EUA. Therefore, the kit should be run on the cobas 6800/8800 platform with cobas 6800/8800 Systems Software v1.2. Performance values listed here are taken from the cobas EUA information. | Testing is limited to the BayCare Laboratories, LLC located at 5455 West Waters Ave. Suite 208, Tampa, FL 33634
PCL SARS-CoV-2 Real-Time RT-PCR Assay | Patients Choice Laboratories, LLC | rRT-PCR | 100% | 100% | nasopharyngeal and oropharyngeal swabs | N gene, S gene, ORF1ab | 10,000 copies/mL | The assay simultaneously detects four targets: three SARS-CoV-2 viral targets, the Nucleocapsid (N) gene, the ORF1ab gene and the Spike protein (S) gene, and one primer/probe set detecting MS2 RNA spiked into the reaction as an extraction and process control. The PCL SARS-CoV-2 Real-Time RT-PCR assay is to be used with the Hamilton Vantage Liquid Handling System (serial number:1630) and the QuantStudio 12 K Flex instrument (software version 1.4). Sensitivity and specificity was determined from 11 positive clinical samples and 19 negative clinical samples. The LoD according to the NIH reference panel is 5400 copies/mL. | Testing is limited to Patients Choice Laboratories located at 7026 Corporate Drive, Indianapolis, IN 46278
Texas Department of State Health Services (DSHS) SARS-CoV-2 Assay | Texas Department of State Health Services, Laboratory Services Section | rRT-PCR | 100% | 100% | nasopharyngeal swabs, oropharyngeal (throat) swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasal aspirates, nasal washes and bronchoalveolar lavage | N gene and ORF1ab | 20 copies/mL | The assay uses primers and probes that were developed and validated under the Emergency Use Authorization (EUA) for the PerkinElmer New Coronavirus Nucleic Acid Detection Kit and which are designed to detect RNA from the SARS-CoV-2 nucleocapsid (N) gene and ORF1ab. The Texas DSHS SARS-CoV-2 Assay is to be usedwith the PerkinElmer Chemagic 360 (software version 6.3.0.3) andThermoFisher Applied Biosystems 7500 Fast Dx Real-Time PCR System (software version 1.4). Sensitivity and specificity was determined from 11 positive clinical samples and 36 negative clinical samples. | Testing is limited to the Texas Department of State Health Services, Laboratory Services Section located at 1100 W. 49th Street, Austin, TX 78756.
DxTerity SARS-CoV-2 RT-PCR Test | DxTerity Diagnostics, Inc | rRT-PCR | 97.2% | 92.5% | saliva | N gene, S gene, ORF1ab | 100 copies/mL | The DxTerity SARS-CoV-2 RT-PCR Test is a real-time reverse transcription polymerase chain reaction (rRT-PCR) test using a MagMax 96 Magnetic Particle Processor (software BindIt 4.0) for RNA extraction and a Thermo Fisher Applied Biosystems ViiA 7 Real-Time PCR amplification system with the Applied Biosystems QuantStudio5 software v1.3 for data analysis. The assay uses primers and probes that were developed and validated under the Emergency Use Authorization (EUA) for the TaqPath COVID-19 Combo Kit. Saliva specimens must be collected using the Spectrum Solutions LLC SDNA-1000 Saliva Collection Device. Saliva specimens must be transported and stored at ambient temperature and tested within 48 hours of collection. Sensitivity and specificity was determined from 36 positive clinical samples and 40 negative clinical samples. | Testing is limited to the DxTerity Diagnostics, Inc. laboratory located at 19500 S. Rancho Way, Suite 116, Rancho Dominquez, California 90220
Guardant-19 | Guardant Health, Inc. | rRT-PCR | 95.5% | 98.3% | nasopharyngeal (NP) swabs, oropharyngeal (OP) swabs, nasal swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasopharyngeal wash/aspirates, nasal aspirates, and nasal washes | N gene (N1 region) | 125 copies/mL | The test uses the same N1 (assay target) and RNase P (endogenous control) primer sequences as those described under the EUA for the CDC 2019-Novel Coronavirus (2019-nCoV) Real-Time RT-PCR Diagnostic Panel to detect the N1 region of the SARS-CoV-2 nucleocapsid (N) gene and the human RNase P gene, respectively. The assay has been optimized for multiple test kits and machines. Sensitivity and specificity was determined from 60 positive clinical samples and 67 negative clinical samples. | Testing is limited to the Guardant Health Clinical Laboratory, 505 Penobscot Drive, Redwood City, CA, 94063
Solaris Multiplex SARS-CoV-2 Assay | Solaris Diagnostics | rRT-PCR | 100% | 100% | (nasopharyngeal swabs, anterior nasal swabs, mid-turbinate nasal swabs, oropharyngeal swabs, nasopharyngeal wash/aspirates, nasal aspirates, bronchoalveolar lavage specimens | N gene (N1 and N2 regions) | 10,000 copies/mL | The Solaris Multiplex SARS-CoV-2 Assay is an rRT-PCR based diagnostic that targets 2 regions of the nucleocapsid gene of SARS-CoV-2. The test should be run on the Applied Biosystems QuantStudio 5 with software version v1.5.1, with RNA extraction performed separately. Sensitivity and specificity were determined from 30 positive and 30 negative clinical samples | Testing is limited to Solaris Diagnostics, located at 110 Dewey Drive Ste A, Nicholasville, KY 40356
Alpha Genomix TaqPath SARS-CoV-2 Combo Assay | Alpha Genomix | rRT-PCR | 96.7% | 100% | nasopharyngeal, oropharyngeal (throat), anterior nasal, and midturbinate nasal swab | ORF1ab, N gene, Spike gene | 4000 copies/mL | This assay uses Taqman technology to target 3 genes in the SARS-CoV-2 genome. This test uses the ThermoFisher Scientific TaqPath COVID-19 Combo Kit, and RNA should be extracted separately. The test should be run on the The Applied Biosystems QuantStudio 12K Flex Real-Time PCR Instrument with QuantStudio RealTime PCR Design and Analysis software version 2.4.0. Sensitivity was determined from 30 positive clinical samples and specificity from 35 negative samples. | Testing is limited to Alpha Genomix Laboratories located at 333 Research Ct, Peachtree Corners, GA 30092
GWU SARS-CoV-2 RT-PCR Test | George Washington University Public Health Laboratory | rRT-PCR | 95% | 100% | nasal, mid-turbinate, nasopharyngeal and oropharyngeal swabs, and BAL | N gene (N1 and N2 regions) | 12500 copies/mL | This is an rRT-PCR assay using the CDC's published N1 and N2 primer and probe sets. RNA extraction should be performed separately, and the test should be run on the Roche LightCycler 480 instrument with Software version 1.5. Sensitivity was determined from 40 clinical samples, and specificity from 42 negative samples. | Testing is limited to GWU Public Health Laboratory, 800 22nd St. NW, Washington, DC 20052
Helix COVID-19 NGS Test | Helix OpCo LLC (dba Helix) | NGS | 100% | 100% | nasopharyngeal swabs, oropharyngeal (throat) swab, mid-turbinate nasal swabs and anterior nasal swabs | S gene | 125 genomic copy equivalents/mL | The Helix COVID-19 NGS Test is a sequencing based assay for detection of SARS-CoV-2 infection. RNA must be extracted separately, and then reverse transcribed and amplified using the TaqPath 1-Step Multiplex Master Mix, run on the Bio-Rad C-1000 Touch thermocycler. The amplified product is then pooled and sequenced. This test uses the Illumina NovaSeq 6000 Sequencing System platform. The assay detects the S gene of SARS-CoV-2, and a human control, RPP30. Sensitivity and specificity were determined from 30 positive and 30 negative clinical samples. | Testing is limited to the Helix Laboratory located at 9875 Towne Centre Drive, San Diego, CA 92121
Poplar SARS-CoV-2 TMA Pooling assay | Poplar Healthcare | transcription mediated amplification (TMA) | 100% | 100% | clinician-instructed self-collected nasal swab specimens, and clinician-collected nasal, nasopharyngeal, and oropharyngeal swab | 2 targets in ORF1ab | 0.01 TCID50/mL | The Poplar SARS-CoV-2 TMA Pooling assay specifically allows for sample pooling of up to 7 samples per run. This assay uses the Aptima SARS-CoV-2 assay, from Hologic. The TMA assay is an isothermal amplification assay using RNA polymerase to generate large pools of replicated RNA. The assay should be run on the Panther platform. Sensitivity was determined from 37 positive clinical samples, and specificity from 32 negative samples. | Testing is limited to Poplar Healthcare, located at 3495 Hack Cross Rd., Memphis, TN 38125
Genus SARS-CoV-2 Assay | ISPM Labs, LLC dba Capstone Healthcare | rRT-PCR | 100% | 100% | nasopharyngeal and oropharyngeal swabs | N gene (2 targets) | 40,000 copies/mL | The Genus SARS-CoV-2 Assay is an rRT-PCR test. This test is run in a 384 well plate, which allows for greater throughput. RNA should be extracted separately. It should be run on the Applied Biosystems QuantStudio 12K Flex Real Time PCR Instrument with QuantStudio RealTime PCR System software version 1.3. Sensitivity and specificity were determined from 43 positive and 28 negative clinical samples. | Testing is limited to the Clinical Virology Laboratory at Capstone Healthcare, located at 8601 Dunwoody Pl. Ste. 444, Sandy Springs, GA 20250
UCSD RC SARS-CoV-2 Assay | University of California San Diego Health | rRT-PCR | 97.7% | 98.3% | nasal, nasopharyngeal, mid-turbinate, anterior nares or oropharyngeal swab specimens | ORF1, E gene | 0.009 TCID50/mL * did not state the TCID50 concentration | The UCSD RC SARS-CoV-2 Assay is a real-time reverse transcription polymerase chain reaction (RT-PCR) test for pooled samples containing up to five individual upper respiratory swab specimens. The assay is the cobas SARS-CoV-2 which has Emergency Use Authorization (EUA) and is designed to detect RNA from SARS-CoV-2 in respiratory specimens from patients as recommended for testing by public health authority guidelines. | The UCSD RC SARS-CoV-2 Assay will be performed at the University of California San Diego Health located at Center for Advanced Laboratory Medicine.
CRL Rapid Response | Clinical Reference Laboratory, Inc. | rRT-PCR | 100% | 100% | Self-collected saliva specimens | RdRP gene | 250 copies/mL | The CRL Rapid Response test is a single-step real-time reverse transcription PCR test based on the Logix Smart Coronavirus Disease 2019 (COVID-19) test (EUA200049). The CoPrimers and probes in the CRL Rapid Response test are designed to detect nucleic acid sequences in the RdRp gene. The CRL Rapid Response test is to be used with the BioRad CFX-96 TOUCH Thermal Cycler instrument using the CFX Manager 3.1 software. Sensitivity and specificity were determined from 32 positive and 27 negative clinical samples. | The CRL Rapid Response test will be performed at Clinical Reference Laboratory, Inc. in Lenexa, KS
SNL-NM 2019 nCoV Real-Time RT-PCR Diagnostic Assay | Sandia National Laboratories | rRT-PCR | 100% | 100% | nasopharyngeal swab, anterior nasal swab, mid-turbinate nasal swab and oropharyngeal swab, nasal washes, nasal aspirates and bronchoalveolar lavage | N gene (N1 and N2 regions) | 6250 copies/mL | The test uses the same primer and probe sets as published by the CDC. The test should be run on the Applied Biosystems 7500 (ABI7500), Applied Biosystems Quant Studio 5, Bio-Rad CFX Connect, or Bio-Rad CFX 96 Real PCR platforms. Sensitivity and specificity were determined from 35 positive and 35 negative clinical samples. | Testing is limited to will be the Sandia National Laboratories’ laboratories located at 1515 Eubank Blvd. SE, Albuquerque, NM 87123 and 7011 East Ave., Livermore, CA 94551,
OraRisk COVID-19 RT-PCR | Access Genetics, LLC | rRT-PCR | 100% (NP swab), 100% (nasal swab) | 100% (NP swab), 95.7% (nasal swab) | nasopharyngeal swab and nasal swab specimens collected in universal transport media, and nasal swabs collected in oral saline rinse | RdRP gene | 15,000 copies/mL | This test uses the Smart Logix Coronavirus 2019 assay, which was approved on April 3, 2020. This test should be run on the LightCycler 480 II instrument with Sequence Detection Software version 1.5.1.62. Nasopharyngeal swab sensitivity and specificity were determined from 15 positive and 15 negative clinical samples. Nasal swab sensitivity and specificity were determined from 22 positive and 23 negative clinical samples. | Testing is limited to the Access Genetics, LLC laboratory, located at 7400 Flying Cloud Drive, Eden Prairie, MN 55344
Compass Laboratory Services SARS-CoV2 Assay | Compass Laboratory Services, LLC | rRT-PCR | 100% | 96% | nasopharyngeal swabs, oropharyngeal (throat) swab, mid-turbinate nasal swabs, anterior nasal swabs, nasopharyngeal wash/aspirate or nasal aspirate from a HCW, or nasal swab from at-home collection | ORF1ab | 5000 copies/mL | The test is in two different wells, one testing for the viral gene and one for the human RNAseP control. It also uses the Primerdesign Ltd Novel Coronavirus Strain 2019-nCoV (RUO) (Ref:Z-Path-2019-nCoV). It can be run on the QuantStudio 12K Flex and QuantStudio 7 Flex platforms. Sensitivity was determined using 24 true clinical samples and specificity from 25 negative clinical samples. | Testing is limited to Compass Laboratory Services, LLC’s laboratory located at 1910 Nonconnah Blvd, Suite 108, Memphis TN 38132
BMC-CReM COVID-19 Test | Boston Medical Center | rRT-PCR | 100% | 100% | nasopharyngeal, oropharyngeal, anterior nasal and mid-turbinate nasal swabs, nasopharyngeal washes/aspirates or nasal aspirates, and bronchoalveolar lavage | N gene (N1 and N2 region) | 10,000 copies/mL (2 step) and 1000 copies/mL (1-step) | The test can be run as a one- or two-step assay. The one-step assay should be run on the QuantStudio 6 Flex platform. The two step assay has cDNA synthesis on the Mastercycler Nexus X2 platform, followed by Taq based amplification. Sensitivity was determined from 30 clinical samples, and specificity from 30 negative samples. | Testing is limited to the Boston Medical Center/Department of Pathology and Laboratory Medicine located at 1 Boston Medical Center Place, Boston, MA 02118
SARS-CoV-2 RNA DETECTR Assay | UCSF Health Clinical Laboratories, UCSF Clinical Labs at China Basin | RT-LAMP | 95% | 100% | nasopharyngeal swabs, oropharyngeal swabs, mid-turbinate nasal swabs, anterior nasal swabs, nasopharyngeal wash/aspirate or nasal aspirate | N gene (N2 region) | 20000 copies/mL | The assay, SARS-CoV-2 RNA DNA Endonuclease-Targeted CRISPR Trans Reporter (DETECTR), performs simultaneous reverse transcription and isothermal amplification using loop-mediated amplification (RT-LAMP). This test uses primer sets and CRISPR guide RNAs (gRNAs) specific for the detection SARS-CoV-2 using reverse transcription-loop-mediated isothermal amplification (RTLAMP) with the Bst DNA polymerase and Cas12-mediated cleavage of the amplified product. The target region of the N gene overlaps with the N2 region of the FDA Emergency Use Authorization (EUA)-authorized CDC assay. The SARS-CoV-2 RNA DETECTR Assay test is to be used with the Thermo Scientific ABI 7500 Fast Dx Real-Time PCR System, Software Version 1.4.1. | Testing is limited to the UCSF Clinical Labs at China Basin in San Francisco, CA.
Gene By Gene SARS-CoV-2 Detection Test | Gene by Gene | rRT-PCR | 100% | 100% | nasal and nasopharyngeal swabs | N gene (N1 and N2 regions) | 6250 copies/mL | The primer and probe sets used are the same as those published by the CDC. RNA extraction should be performed separately, and the test should be run using Bio-Rad Reliance 1-Step Multiplex Supermix on the Bio-Rad CFX384 Touch RT PCR platform. Sensitivity was determined using 35 positive clinical samples, and specificity from 36 negative samples. | Testing is limited to Gene By Gene’s laboratory located at 1445 North Loop West Suite 760, Houston, TX 77008
Laboratorio Clinico Toledo SARS-CoV-2 Assay | Laboratorio Clinico Toledo | rRT-PCR | 100% | 100% | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swab specimens and BAL specimens | E gene, N gene, RdRP | 203 copies/mL | The SARS CoV-2 Assay is a real-time reverse transcription polymerase chain reaction (RT-PCR) test. The test uses one primer and probe to detect one region in the SARS CoV-2 nucleocapsid N gene, one primer and probe for the detection of SARS like coronavirus E gene and one primer and probe to detect RNAase P (RdRP). This test uses the Roche Magna Pure 96 system (Software version 3.1.1) for the extraction process and the Roche Light Cycler 480 (software version LCS480 1.5.16.2 SP@-UDF v 2.0) for the RT-PCR amplification process. | Testing is limited to Laboratorio Clinico Toledo in Arecibo, Puerto Rico.
CentoFast-SARS-CoV-2 RT-PCR Assay | Centogene US, LLC | rRT-PCR | 100% | 100% | dry oropharyngeal swabs | E gene and RdRP | 5000 copies/mL | The device tests in a tiered approach in which samples positive for the E-gene template are followed up with a confirmatory testing for E and RdRP targets. The E-target is the same in both tiers. The CentoFast-SARS-CoV-2 Assay is to be used with KingFisher Flex (Thermo Fisher) for RNA extraction and the LightCycler 480 II (Roche) for the RT-PCR. The LightCycler 480 II is using Software Version 1.5.1.62 or higher. | Testing is limited to CENTROGENE, LLC.
LifeHope 2019-nCoV Real-Time RT-PCR Diagnostic Panel | LifeHopes Lab | rRT-PCR | 100% | 100% | nasal, mid-turbinate, nasopharyngeal, oropharyngeal swab specimens, and bronchoalveolar lavage specimens | N gene (N1 and N2 regions) | 2500 copies/mL | Testing is performed on the QuantStudio 5 (ThermoFisher) and analysis is done using the QuantStudio Design and Analysis Software version 1.5.1. The SARS-CoV-2 primer and probe set(s) provided by Integrated DNA Technologies (IDT) are designed to detect RNA from the SARS-CoV-2 in upper respiratory specimens and bronchoalveolar lavage from patients. All nucleic acid is extracted using the IndiMag Extraction Kit (Indical, SP54106 IB). Sensitivity and specificity were determined from 30 positive and 30 negative clinical samples. | Testing is limited to the LifeHope Labs located in Sandy Springs, GA.
Inform Diagnostics SARS-CoV-2 RT-PCR Assay | Inform Diagnostics, Inc. | rRT-PCR | 93.9% | 100% | nasopharyngeal, oropharyngeal, anterior nasal, and mid-turbinate nasal swabs, as well as nasopharyngeal wash/aspirate or nasal aspirates, and bronchoalveolar lavage | N gene (N1 and N2 regions) | 20,000 copies/mL | The test uses the same primer and probe sets as published by the CDC. The test should be run on the e Applied Biosystems 7500 Fast Dx platform with software version 1.4 or the QuantStudio with software version 1.0.3. RNA extraction should be performed separately. Sensitivity was determined from 33 positive clinical samples, and specificity from 31 samples. | Testing is limited to Inform Diagnostics, Inc. located in Phoenix, AZ
DSL COVID-19 Assay | Diagnostic Solutions Laboratory, LLC | rRT-PCR | 100% | 97% | nasal, midturbinate, nasopharyngeal, and oropharyngeal swabs and bronchioalveolar lavage | N1 and N3 regions of the N gene, and the spike gene | 10 copies/swab | The N1 and N3 primers are from the published set by the CDC, while Diagnostic Solutions designed the Spike protein primers and probe. RNA extraction should be performed separately, and the test should be QuantStudio 5 Real-Time PCR System (QS5) with QS5 Analysis software. Sensitivity and specificity were determined from 32 positive and 32 negative clinical samples. | Testing is limited to the Diagnostic Solutions Laboratory, LLC in Alpharetta, GA
MD Anderson High-throughput SARS-CoV-2 RT-PCR Assay | University of Texas MD Anderson Cancer Center, Molecular Diagnostics Laboratory | rRT-PCR | 100% | 100% | nasopharyngeal, oropharyngeal, nasal, and mid-turbinate swabs | N gene (N1 and N2 regions) | 5000 copies/mL | The test uses the same primer and probes as published by the CDC, as well as an internal human control gene of RNAseP. They publish 3 different processing pipelines, with different RNA extraction procedures (manual vs automated) and different sample aliquoting techniques. RNA extraction should be performed separately, and the test should be run on QuantStudio 7 Flex Fast Sequence Detection System (Software V1.3), Applied Biosystems 7500 (Software V1.5.1), or QuantStudio Dx Real-Time PCR , 96 well fast block (Software V1.0.3). The sensitivity was determined from 31 positive clinical samples, and specificity from 28 negative samples. | Testing is limited to the University of Texas MD Anderson Cancer Center in Houston, Texas.
HealthQuest Esoterics TaqPath SARS-CoV-2 Assay | HealthQuest Esoterics | rRT-PCR | 100% | 100% | nasopharyngeal, oropharyngeal, anterior nasal, and mid-turbinate nasal swabs, as well as nasopharyngeal wash/aspirate or nasal aspirates, and bronchoalveolar lavage | N, S, and ORF1ab | 20,000 copies/mL | The test uses the same primers and probes as the ThermoFisher TaqPath COVID-19 Combo Kit. The test should be run on the QuantStudio 12K Flex Real-Time PCR System. Sensitivity was determined from 35 positive clinical samples, and specificity from 34 negative samples. The LoD according to the NIH reference panel is 18,000 NDU/mL. | Testing is limited to HealthQuest Esoterics
FRL SARS CoV-2 Test | UAB Fungal Reference Lab | rRT-PCR | 100% | 100% | nasopharyngeal swab, oropharyngeal swab, nasal swab, mid-turbinate nasal swab, anterior nasal swab, nasopharyngeal wash/aspirate, nasal aspirate, and bronchoalveolar lavage specimens | N1 region of N gene | 125 copies/mL | This test uses the CDC N1 gene primer, but not the N2. It should be run on the QuantStudio 5 platform. Sensitivity and specificity were determined from 30 positive and 30 negative clinical samples. | Testing is limited to the UAB Fungal Reference Lab in Birmingham, Alabama
OSUWMC COVID-19 RT-PCR test | The Ohio State University Wexner Medical Center | rRT-PCR | 100% | 100% | nasal, oropharyngeal and nasopharyngeal swab | N gene (N1 and N2 regions) | 250 copies/mL | The primer and probe sequences are identical to those listed by the CDC. This test should be run on the Applied Biosystems 7500 platform, with RNA extraction performed separately. Sensitivity and specificity listed here were determined using 26 positive and 25 negative true clinical samples. | Testing is limited to the Ohio State University Wexner Medical Center.
Omni COVID-19 Assay by RT-PCR | Omnipathology Solutions Medical Corporation | rRT-PCR |  |  | nasopharyngeal swab, oropharyngeal swab, anterior nasal swab, mid-turbinate nasal swab, nasopharyngeal wash/aspirate, nasal aspirate, and bronchoalveolar lavage | N gene (N1 and N2 regions) | 1230 copies/mL | This is an all-in-one test, with RNA extraction and rRT-PCR completed.This test can be run on the BD Max system, and can process 110 tests per run.The primers and probes are the same as listed by the CDC. Sensitivity and specificity listed by the manufacturer are only from contrived clinical samples, not true clinical samples at this time. | Testing is limited to the Omnipathology Solutions Medical Corporation in Pasadena, CA.
SARS-CoV-2 RT-PCR assay | Cormeum Laboratory Services | rRT-PCR | 100% | 100% | nasopharyngeal swabs | N gene, S gene, ORF1ab | 200 copies/mL | The SARS-CoV-2 primer and probe set designed to amplify and detect three regions of the SARS-CoV-2 single stranded RNA genome: the Nucleocapsid (N), Spike (S) and Open Reading Frame 1ab (Orf1ab) in respiratory specimens from patients as recommended for testing by public health authority guidelines. Sensitivity and specificity were determined from 37 positive samples and 147 negative samples. | Healthcare providers can send clinical specimens to Cormeum Laboratory Services
Warrior Diagnostics SARS-CoV-2 Assay | Warrior Diagnostics, Inc. | rRT-PCR | 100% | 100% | nasopharyngeal swabs | N gene (N1 and N2 region) | 150,000 copies/mL | The test uses the same primer and probes as the published CDC set, and should be run on the Agilent AriaMx Real-Time PCR System. RNA extraction must be performed separately. Up to 30 samples can be run at once. The LOD listed is using plasmids with SARS-CoV-2 genes inserted, as the company did not have access to SARS-CoV-2 RNA at the time. Sensitivity and specificity was determined from 31 positive samples and 29 negative samples. | Healthcare providers can send clinical specimens to Warrior Diagnostics, Inc
NEcov19 RT-PCR Assay | Nebraska Medicine Clinical Laboratory | rRT-PCR | 100% | 100% | nasopharyngeal swabs, oropharyngeal swabs, anterior nasal swabs, mid-turbinate swabs, nasal washes, nasal aspirates and bronchoalveolar lavage (BAL) | E gene | 1000 copies/mL | The assay detects a nucleic acid sequence from the SARS-CoV-2 E gene. The RTPCR assays are performed on the Applied Biosystems QuantStudioDx RT-PCR instrument with QuantStudio Dx with QuantStudio Test Development Software version 1.0.3 | Healthcare providers can send clinical specimens to the Nebraska Medicine Clinical Laboratory.
Aspirus SARS-CoV-2 rRT-PCR Assay | Aspirus Reference Laboratory | rRT-PCR | 100% | 100% | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swab specimens, and bronchoalveolar lavage | N gene (N1 and N2 region) | 500 copies/mL | The test should be performed on the BD Max platform. The test uses primer and probe sequences, including human internal controls, published by the CDC. The RNA extraction is performed as part of the test, using the BD MAX TNA-3 kit. Sensitivity and specificity were determined using 30 positive and 30 negative clinical samples. | Healthcare providers can send clinical specimens to Aspirus Laboratory.
CSI SARS-CoV-2 RT PCR Test | CSI Laboratories | rRT-PCR | 100% | 100% | nasopharyngeal swabs, nasal, and oropharyngeal swabs, and bronchoalveolar lavage samples | N gene (N1 and N2 region) | 6250 copies/mL | The test should be performed using the Applied Biosystem Quantstudio 7 instrument (software version 1.3), and is an all-in-one rRT-qPCR reaction. RNA extraction must be performed separately. The primers and probes (including the control human gene) are from the CDC published sequences. Sensitivity and specificity were determined from 25 positive and 30 negative clinical samples. | Healthcare providers can send clinical specimens to CSI laboratories
SARS-CoV-2 (N gene detection) Test | Exact Sciences Laboratories | rRT-PCR | 95% | 100% | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swab specimens | N gene (two targets) | 2600 copies/mL | The test uses specific primer and probe sets developed by CDC to detect two regions (nCOV_N1 and nCov_N2) in the SARS-CoV-2 nucleocapsid (N) gene and one primer and probe set to detect human RNase P (RP) in a clinical sample. This test uses the ABI 7500 Fast Dx (Applied Biosystems) instrument with Sequence Detection Software (SDS) version 1.4.1. Sensitivity and specificity were determined from 40 positive and 38 negative clinical samples. | Healthcare providers can send clinical specimens for testing to the Exact Sciences Laboratories laboratory
Express Gene 2019-nCoV RT-PCR Diagnostic Panel | Express Gene LLC, DBA: Express Gene Molecular Diagnostics Laboratory | rRT-PCR | 100% | 100% | nasopharyngeal swabs, oropharyngeal swabs, anterior nasal swabs, mid-turbinate nasal swabs, nasopharyngeal aspirate/wash or nasal aspirates, and BAL | N gene, S gene, ORF1ab | 800 copies/mL for the S gene, 320 copies/mL for the ORF1ab region, and 160 copies/mL for the N gene | The test detects three specific regions of the SARS-CoV-2 genome including the ORF1ab region and the N (nucleocapsid) and S (Spike protein) genes. RNA is reverse transcribed to cDNA using the TaqPath 1-Step Multiplex Master Mix and subsequently amplified using the QuantStudio 12K Flex instrument with software version 2.2.3. Sensitivity and specificity were determined from 30 positive and 30 negative clinical samples. | Healthcare providers can send clinical specimens for testing to the Express Gene LLC dba Molecular Diagnostics Laboratory
Avera Institute for Human Genetics SARS-CoV-2 Assay | Avera Institute for Human Genetics | rRT-PCR | 100% | 100% | nasopharyngeal, nasal, and oropharyngeal swab specimens, and BAL | N gene (two targets) | 1600 copies/mL | The assay is a real-time reverse transcription polymerase chain reaction (rRT -PCR) test based on the EUA FDA issued for the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel. The test is to be used with the Applied Biosystems Quant Studio 7 Flex equipped with Quant Studio software v1.1, or the Applied Biosystems Viia 7 equipped with the Quant Studio software v1.1. Sensitivity and specificity were determined from 30 positive and 30 negative clinical samples. | Healthcare providers can send clinical specimens for testing to the Avera Institute for Human Genetics laboratory
One Health Laboratories SARS-CoV-2 Real-Time RT-PCR Test | One Health Laboratories, LLC | rRT-PCR | 100% | 97.8% | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swab specimens, and nasopharyngeal wash/aspirate or nasal aspirate specimens | ORF1ab, spike protein | 1800 copies/mL | Test reagents, including extraction kits, are from Biomeme. The thermocycler is the Biomeme Franklin three9 Real-Time PCR Thermocycler, which is run through a smartphone app. The test has an MS2 bacteriophage internal control that has a third primer and probe. The sensitivity and specificity were determined using contrived samples, but the values reported here are from 76 clinical samples (35 positive, 45 negative, 1 inconclusive). | Healthcare providers can send clinical specimens for testing to OneHealth Laboratories.
Cedars-Sinai SARS-CoV-2 Assay | Cedars-Sinai Medical Center, Department of Pathology and Laboratory Medicine | rRT-PCR | 100% | 100% | nasopharyngeal swabs | NSP-3 gene | 10,000 copies/mL | RNA should be isolated using an appropriate kit, such as the Qiagen QIA Connect or QIAsymphony. The test uses the A* STAR Fortitude kit 2.0, and should be run on the AB QuantStudio 6 Flex (v1.3) or the ABI 7500 Fast Real Time PCR(v1.4). Sensitivity and specificity were determined using 30 positive and 30 negative clinical specimens (as measured by 3 different EUA approved tests). | Healthcare providers can send clinical specimens for testing to Cedars-Sinai Medical Center, Department of Pathology and Laboratory Medicine.
Triplex CII-SARS-CoV-2 rRT-PCR Test | Columbia University Laboratory of Personalized Genomic Medicine | rRT-PCR | 100% | 100% | nasopharyngeal, oropharyngeal, nasal, and mid-turbinate nasal swab | Nucleoprotein (two targets - N1 and N2), human RP | 280 copies/mL | This test is called the Triplex CII-SARS-CoV-2 rRT-PCR test, which should be run with the bioMérieux easyMAG automated extraction platform/NucliSENS easyMAG extraction kit and the ABI Fast with software version 2.3. Sensitivity and specificity were determined using contrived samples, but also 17 positive and 23 negative clinical samples. The values presented here are representative of the clinical (not contrived) samples. | Healthcare providers can send clinical specimens for testing to Columbia University Laboratory of Personalized Genomic Medicine
Biocollections Worldwide SARS-Co-V-2 Assay | Biocollections Worldwide, Inc. | rRT-PCR | 100% | 100% | nasal swabs, nasopharyngeal swabs, oropharyngeal swabs, nasopharyngeal wash/aspirate or nasal aspirate, BAL | Nucleoprotein (three targets - N1, N2, N3), human RP | 1000 copies/mL | The SARS-CoV-2 Assay is to be used with the Abbott automated m2000 System which combines the automated sample preparation system, m2000sp (software version 8.1.8.0) and the real-time PCR thermal cycler/reader system, m2000rt (software version 8.1.9.0). | Healthcare providers can send clinical specimens for testing to Biocollections Worldwide Inc., laboratory
UTHSC/UCH SARS-CoV-2-RT-PCR Assay | UTMG Pathology Laboratory | rRT-PCR |  |  | nasal swabs | Nucleoprotein (two targets - N1 and N2), human RP | 100 plaque forming units (PFU) | The UTHSC/UCH SARS-CoV-2 assay is real time RT-PCR test, which uses real time PCR primer and fluorescent probe based chemistry with CDC developed EUA primer/probe sets purchased from Integrated DNA Technologies (IDT) to detect two regions of SARS-CoV-2 nucleocapsid (NI & N2) gene and a third primer/probe set to detect human RNase P (RP) in clinical samples, in conjunction with reverse transcription and PCR using Applied Biosystems TaqPath 1-Step RT-qPCR Master Mix, CG (ThermoFisher Catalog # A15299). | Healthcare providers can send clinical specimens for testing to UTMG Pathology Laboratory
Altru Dx SARS-CoV-2 Assay | Altru Diagnostics, Inc. | rRT-PCR | 100% | 100% | nasal, midturbinate, nasopharyngeal, and oropharyngeal swab specimens | Nucleoprotein, spike, ORF1ab | 625 copies/mL | The Altru Diagnostics SARS-CoV-2 Test uses the Thermo Fisher Scientific TaqMan 2019-nCoV Assay Kit v1 (singleplex) RUO Kit. 11 clinical specimens were used to calculate sensitivity. | Healthcare providers can send clinical specimens for testing to Altru Diagnostics, Inc.
SAR-Co-V-2 Assay | Biocerna | rRT-PCR | 100% for all targets (on Copan swab) On Genotek swab: 100% for N gene, 95% for ORF1ab, 85% for S gene | 100% for all targets | nasal, mid-turbinate, nasopharyngeal, and oropharyngeal swab and bronchoalveolar lavage (BAL) | Nucleoprotein, spike, ORF1ab | 250 copies/mL (Genotek swabs) 375 copies/mL (Copan swabs) | RT-PCR including cDNA synthesis and PCR amplification of the target sequences is performed on the QuantStudio 5 real-time PCR system. Specimens can be collected using the OraCollect-RNA (OR-100) swabs (Genotek), Copan eSwab (Copan), or other swabs systems found to be equivalent. Maximum 20 specimens were used to calculate sensitivity | Healthcare providers can send clinical specimens for testing to Biocerna LLC.
SARS-CoV-2 Assay | Nationwide Children’s Hospital | rRT-PCR | 100% | 100% | nasopharyngeal swabs, nasal swabs, mid-turbinate nasal swabs, oropharyngeal swabs, nasopharyngeal wash/aspirates, nasal aspirates, bronchoalveolar lavage, lower respiratory tract aspirates, and sputum | Nucleoprotein (two targets - N1 and N2), human RP | 15 copies/rxn (rxn volume of 20 ul; for nasopharyngeal swabs) 200 copies/uL (for sputum samples) | The NCH SARS-CoV-2 assay is a modification of the CDC 2019-nCoV Real-Time RT-PCR assay which utilizes the same SARS-CoV-2 N1 and N2, and human RNAase P control primers and probes for target amplification and detection. LOD was calcuated with 10 specimens | Healthcare providers can send clinical specimens for testing to Nationwide Children’s Hospital in Columbus, Ohio
UDX SARS-CoV-2 Molecular Assay | Ultimate Dx Laboratory | rRT-PCR | 100% | 100% | nasal swabs, nasopharyngeal swabs, oropharyngeal swabs | ORF1ab | 100 copies/mL | The UDX SARS-CoV-2 Molecular Assay is to be used with the Applied Biosystems QuantStudio 12K platform (QuantStudio 12K Flex Software v1.4). The test is intended to be used with the automated extraction procedure KingFisher Flex Purification System (ThermoFisher Scientific, Cat#501527924). | Healthcare providers can send clinical specimens for testing to Ultimate Dx Laboratory
SARS-CoV-2 Assay | AIT Laboratories | rRT-PCR | 100% | 100% | nasal, midturbinate, nasopharyngeal and oropharyngeal swab specimens, sputum, BAL | Nucleoprotein, spike, ORF1ab | 1000 copies/mL | The SARS-CoV-2 Test uses the ThermoFisher Scientific TaqPath COVID-19 Combo Kit that was FDA authorized for emergency use (EUA) on March 13, 2020. RNA isolated from respiratory specimens is reverse transcribed to cDNA and subsequently amplified using the QuantStudio 12K Flex real-time PCR system (Applied Biosystems) with Software version 1.5.1 | Healthcare providers can send clinical specimens for testing to the AIT Laboratories
COVID-19 Key | Southwest Regional PCR Laboratory LLC. dba MicroGen DX | rRT-PCR | 100% | 100% | nasopharyngeal, oropharyngeal, anterior nasal, and mid-turbinate nasal swabs, nasopharyngeal awash/aspirate or nasal aspirates, sputum | Nucleoprotein (three targets - N1, N2, N3), human RP | 100 copies/mL | The test includes the same N1 and N2 primer and probe sequences as the CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel for the specific detection of the 2019-nCoV, and the N3 primer and probe sequences for the universal detection of SARSlike coronaviruses. Extracted RNA is reverse transcribed to cDNA and subsequently amplified using the Roche Lightcycler 480II with Lightcycler 480 software version 1.5.1 sp3 | Healthcare providers can send clinical specimens for testing to MicroGen DX
SARS-CoV-2 PCR Test | Diatherix Eurofins Laboratory | rRT-PCR | 98.9% (at 2x LOD) | 100% | nasal swabs, nasopharyngeal swabs/wash, oropharyngeal swabs/throat swabs, bronchial aspirates, or sputum | Spike gene | 1000 copies/mL | Diatherix’s Target-Enriched Polymerase Chain Reaction is a nested, end-point PCR technology. Inside nested primers have a unique tag sequence complementary to proprietary SuperPrimers (Fs and Rs), which are included in the primer mix. The universal SuperPrimers are used to amplify all targets by annealing to the complementary tag sequence on the inside nested primers. Sensitivity was calculated with a maximum of 88 clinical specimens | Healthcare providers can send clinical specimens for testing to Diatherix Eurofins Laboratory
SARS-CoV-2, Molecular Detection Assay | MayoClinic | rRT-PCR | 100% | 100% | nasopharyngeal swabs, throat swabs, nasal swabs, sputum, tracheal secretions, BAL fluid, and bronchial washings | Nucleoprotein, ORF1ab | 156 copies/mL | RNA is isolated/purified from upper and lower respiratory specimens using the bioMérieux easyMAG and EMAG and is reverse transcribed to cDNA and subsequently amplified using Roche LightCycler 480. | Healthcare providers can send clinical specimens for testing to Mayo Clinic Laboratories, Rochester, MN
CDI Enhanced COVID-19 Test | HUMC Molecular Pathology Laboratory | rRT-PCR | 100% at 2x LOD | 100% | nasopharyngeal, oropharyngeal, anterior nasal, and mid-turbinate nasal swabs; nasopharyngeal wash/aspirate or nasal aspirate | Nucleoprotein (N2 target); E gene (to detect all CoVs); human RP | 4000 copies/mL | The CDI Enhanced COVID-19 Test is a real-time reverse transcription polymerase chain reaction (RT-PCR) test. The CDI Enhanced COVID-19 Test is to be used with the MagNA Pure 24 System and the Mic PCR cycler with micPCR software version2.8.0 or version2.8.13. | Healthcare providers can send clinical specimens for testing to Hackensack University Medical Center (HUMC) Molecular Pathology Laboratory
CirrusDx SARS-CoV-2 Assay | CirrusDx Laboratories | rRT-PCR | 95% | 100% | nasal swabs, nasopharyngeal swabs, oropharyngeal swabs, and BAL | Nucleoprotein, E gene, RdRP | 78 copies/rxn | RNA from nasal swabs, nasopharyngeal swabs, oropharyngeal swabs and BAL specimens extracted using the Qiagen QIAamp Viral Mini Kit or the KingFisher Flex automated extraction is reverse transcribed to cDNA and subsequently amplified using the QuantStudioTM 5 or 7. | Healthcare providers can send clinical specimens for testing to CirrusDx Laboratories
Childrens-Altona-SARS-CoV-2 Assay | Infectious Diseases Diagnostics Laboratory (IDDL), Boston Children’s Hospital | rRT-PCR | 100% | 100% | nasopharyngeal, oropharyngeal, anterior nasal, and mid-turbinate nasal swab | Spike gene; E gene (to detect all CoVs) | 2300 copies/mL | RNA is isolated from upper respiratory specimens using the Qiagen Virus MinElute Spin kit on the QiaCube extraction platform and is reverse transcribed to cDNA and subsequently amplified using the Qiagen Rotor-Gene Q 5Plex RT-PCR Platform with software version 2.3.1 | Healthcare providers can send clinical specimens for testing to Infectious Diseases Diagnostics Laboratory (IDDL), Boston Children's Hospital
SARS-CoV-2 Test | Exact Sciences Laboratories | rRT-PCR | 100% | 100% | nasal, nasopharyngeal, and oropharyngeal swab | Nucleoprotein, E gene, RdRP | 600 copies/mL | RNA isolated from upper respiratory specimens is reverse transcribed to cDNA and subsequently amplified using the QuantStudio 5 (QS5) realtime PCR system (Applied Biosystems) with Design and Analysis Software version 1.5.1. | Healthcare providers can send clinical specimens for testing to Exact Sciences Laboratories
SARS-CoV-2 Assay | Integrity Laboratories | rRT-PCR | 100% | 100% | nasal, nasopharyngeal and oropharyngeal swab | Nucleoprotein (two targets - N1 and N2), human RP | 10000 copies/mL | The Integrity Laboratories SARS-CoV-2 RT-PCR assay for SARS-CoV-2 (originally called 2019 Novel Coronavirus or 2019 nCoV) testing utilizes the CDC-developed assay that targets the Nucleocapsid gene of this virus. Two gene targets, N1 and N2, are used to detect cases of COVID-19. The SARS-CoV-2 RT-PCR test is to be used with the TaqPath 1-step RT-qPCR Master Mix (ThermFisher Scientific) on the Applied Biosystems QuantStudio 12K Flex RealTime PCR System. Results are analyzed with the QuantStudio 12K Flex Software v1.3. | Healthcare providers can send clinical specimens for testing to Integrity Laboratories
COVID-19 RT-PCR Test | Pathology/Laboratory Medicine Lab of Baptist Hospital, Miami, Florida | rRT-PCR | 93.3% | 96.7% | nasopharyngeal and oropharyngeal swabs | Nucleoprotein (two targets - N1 and N2), human RP | 500 copies/mL | SARS-CoV-2 nucleic acid from patient samples is extracted utilizing the Promega Maxwell RSC Viral Total Nucleic Acid Purification Kit and then reverse transcribed to complementary DNA (cDNA) which is then amplified during the PCR reaction using the ThermoFisher Scientific TaqPath 1-Step RT-qPCR Master Mix. The COVID-19 RT-PCR test is to be used with the Applied Biosystems QuantStudio 7 Flex using QuantStudio Real-Time PCR software. The following extraction platform can be used: Promega Maxwell® CSC/RSC Instrument. | Healthcare providers can send clinical specimens for testing to the Pathology/Laboratory Medicine Lab of Baptist Hospital Miami
Orig3n 2019 Novel Coronavirus (COVID-19) Test | Orig3n, Inc. Laboratory | rRT-PCR |  |  | oropharyngeal, nasopharyngeal, anterior nasal, and mid-turbinate nasal swab | Nucleoprotein (three targets - N1, N2, N3), human RP | 10000 copies/mL | RNA is isolated using the Omega Bio-Tek Mag-Bind Viral DNA/RNA 96 bead kit and is reverse transcribed to cDNA and subsequently amplified using the Applied Biosystems QuantStudio7 Flex (QS7) instrument with QuantStudio Real-Time PCR software version 1.3. | Healthcare providers can send clinical specimens for testing to Orig3n, Inc.
SDI SARS-CoV-2 Assay | Specialty Diagnostic (SDI) Laboratories | rRT-PCR | 100% | 100% | nasopharyngeal and oropharyngeal swabs | ORF1ab | 500 copies/mL | The SDI SARS-CoV-2 assay is a real-time reverse transcription polymerase chain reaction (rRT -PCR) test. The primer and probe set, used in this test is identical to that described in the BGI emergency use authorized assay (EUA200034). The SDI SARS-CoV-2 assay is to be used with AusDx MT-Prep Extraction System and Applied Biosystems 7500 instrument with software version 2.0.5. | Healthcare providers can send clinical specimens for testing to Specialty Diagnostic (SDI) Laboratories
UNC Health SARS-CoV-2 real-time RT-PCR test | University of North Carolina Medical Center | rRT-PCR | 100% | 100% | nasopharyngeal and oropharyngeal specimens, tracheal aspirates, BAL/bronchial washings (BW), and nasopharyngeal aspirates (NA)/nasal washings (NW) | E gene | 1000 copies/mL | This real-time RT-PCR test detects SARS-CoV-2 by targeting the E gene. The UNC SARS-CoV-2 real-time RT-PCR test is to be used with the bioMerieux EasyMag or Roche MagNAPure 24 and the ABI PRISM 7500 instruments. | Healthcare providers can send clinical specimens for testing to University of North Carolina Medical Center
MGH COVID-19 qPCR assay | Massachusetts General Hospital | rRT-PCR | 100% | 100% | nasal, nasopharyngeal and oropharyngeal swab and BAL | Nucleoprotein (two targets - N1 and N2), human RP | 5000 copies/mL | The MGH SARS-CoV-2 for SARS-CoV-2 (originally called 2019 Novel Coronavirus or 2019 nCoV) testing utilizes the CDC-developed assay that targets the Nucleocapsid gene of this virus. RNA extraction is conducted using the QIAamp Viral RNA Mini Kit (Qiagen). The SARS-CoV-2 RT-PCR test is to be used with the TaqPath 1-step RT-qPCR Master Mix (ThermFisher Scientific) on the following validated instruments: (1) Cobas Z 480 real-time PCR analyzer and analysis software (Roche), (2) QuantStudio 7 Flex Real-Time PCR System and analysis software (ThermoFisher), (3) ViiA 7 Real-Time PCR System and analysis software (ThermoFisher) | Healthcare providers can send clinical specimens for testing to Massachusetts General Hospital
SARS-Cov-2 Assay | Northwestern Diagnostic Molecular Laboratory | rRT-PCR |  |  | nasopharyngeal swab, oropharyngeal swab, nasal swab, mid-turbinate nasal swab and BAL | Nucleoprotein (one target - N1), human RP | 100000 copies/mL | The Northwestern Medicine SARS-CoV-2 Assay is a real-time reverse transcription polymerase chain reaction (rRT-PCR) test that uses commercially available extraction reagents and primers and probes that were developed and validated by the U.S. Centers for Disease Control and Prevention. The Northwestern Medicine SARS-CoV-2 Assay is for use with the ThermoFisher Scientific Quant Studio 6 Flex, equipped with Quant Studio Real-time PCR software v1.1. | Healthcare providers can send clinical specimens for testing to Northwestern Diagnostic Molecular Laboratory
SARS-CoV-2 RT-PCR test | Infectious Disease Diagnostics Laboratory within the Department of Pathology and Laboratory Medicine at the Children’s Hospital of Philadelphia | rRT-PCR | 96.7% | 100% | nasopharyngeal aspirates, nasopharyngeal swabs, nasal swabs, mid-turbinate nasal swabs, tracheal aspirates, and BAL | Nucleoprotein (one target - N2) | 20,747 copies/mL | The SARS-CoV-2 RT-PCR test is to be used with the Roche MagNAPure LC Total Nucleic Acid automated extraction platform and the Applied Biosystems Quant Studio DX automated PCR platform using software version 1.0.1. | Healthcare providers can send clinical specimens for testing to Department of Pathology and Laboratory Medicine at the Children’s Hospital of Philadelphia Clinical Laboratory
SARS-CoV-2 PCR test | Yale New Haven Hospital, Clinical Virology Laboratory | rRT-PCR | 100% | 100% | nasopharyngeal swabs, oropharyngeal swabs, nasal swabs, mid-turbinate nasal swabs, BALs, and sputum | Nucleoprotein (two targets - N1 and N2), human RP | 2000 copies/mL for N1 4000 copies/mL for N2 | SARS-CoV-2 (originally called 2019 Novel Coronavirus or 2019 nCoV) testing utilizes the CDC-developed assay that targets the Nucleocapsid gene of this virus. Two gene targets, N1 and N2, are used to detect cases of COVID-19. RNA extraction is conducted using the NUCLISENS EASYMAG from Biomerieux. The SARS-CoV-2 RT-PCR test is to be used with the TaqPath 1-step RT-qPCR Master Mix on the ABI QuantStudio Flex 6 Real-Time PCR System with software version 1.3 | Healthcare providers can send clinical specimens for testing to the Clinical Virology Laboratory at Yale New Haven Hospital